Federal Government of Somalia

Ministry of Finance of Somalia

Request for Bids

GOODS

Supply, Delivery, Installation and Commissioning of Laboratory Equipment for the National Reference Laboratory (NRL) of the Somali Bureau of Standards (SoBS)

| # IFB No: |  | # SO-MOF-561587-GO-RFB |
| --- | --- | --- |
| # Bid type: |  | # Single Stage – Two Envelope Bidding with Rated Criteria. |
| # Project: |  | # De-Risking, Inclusion and Value Enhancement of Pastoral Economies Project (Drive) |
| # Project ID: |  | # P176517 |
| # Purchaser: |  | # Ministry of Finance, Somalia |

Request for Bids

GOODS

Purchaser: 		Ministry of Finance – Somalia

Project:               De-risking, inclusion and value enhancement of pastoral economies in the horn of Africa (DRIVE)

Contract title:  Supply, Delivery, Installation and Commissioning of Laboratory Equipment for the National Reference Laboratory (NRL) of the Somali Bureau of Standards (SoBS)

Country: 		Federal Government of Somalia

Project ID: 		P176517

RFB No:	 	SO-MOF-561587-GO-RFB

- The Federal Government of Somalia has received a loan from the World Bank toward the cost of The De-risking, inclusion and value enhancement of pastoral economies in the horn of Africa (DRIVE) and it intends to apply part of the proceeds toward payment under the contract for Supply, Delivery, Installation and Commissioning of Laboratory Equipment for the National Reference Laboratory (NRL) of the Somali Bureau of Standards (SoBS). for this contract, the Borrower shall process the payments using the Direct Payment disbursement method, as defined in the World Bank’s Disbursement Guidelines for Investment Project Financing, except for those payments, which the contract provides to be made through letter of credit.

- The Ministry of Finance of the Federal Government of Somalia now invites sealed Bids from eligible Bidders for Supply, Delivery, Installation and Commissioning of Laboratory Equipment for the National Reference Laboratory (NRL) of the Somali Bureau of Standards (SoBS). The Contract includes Related Services comprising site-readiness assessment, delivery logistics and customs support, installation, commissioning, calibration/verification and IQ/OQ/PQ qualification, training, proficiency-testing enrolment, external calibration services, handover documentation, cybersecurity hardening, and warranty support. Bids will be evaluated on the basis of rated criteria.

- Bidding will be conducted using Request for Bids (RFB) selection method with Open International approach as specified in the World Bank’s “Procurement Regulations for IPF Borrowers” September 2023 (“Procurement Regulations”), and is open to all eligible Bidders as defined in the Procurement Regulations that meet the following minimum qualification criteria:

- Bidder’s Financial Capability: The Bidder shall furnish documentary evidence that it meets the following financial requirement(s): minimum average annual turnover in the last three (3) years (2023, 2024 and 2025) of not less than USD 3,000,000 (Three million United States Dollars), and access to liquid assets, lines of credit or other financial resources, net of other contractual commitments, of not less than USD 750,000 (seven hundred and fifty thousand United States Dollars).

- Bidder’s Experience and Technical Capacity: During the past five (5) years, the Bidder must have successfully completed at least two (2) contracts involving goods of similar nature and complexity (supply, delivery, installation and commissioning of laboratory, analytical, scientific or medical equipment), each with a minimum value of USD 2,500,000 (Two Million Five Hundred Thousand Dollars).

- Bidder’s Legal Requirements: The Bidder shall be legally registered in its country of origin.

- Interested eligible Bidders may obtain further information from Ministry of Finance of Somalia, and inspect the bidding document during office hours 10 AM to 4 PM Somalia local time at the address given below. A pre-bid meeting will be held on 12th November 2026 at 11:00 AM. Somalia local time. Participation details are available on request from the address below.

- A complete set of bidding documents in English language will be provided upon request, free of charge.  The document including any subsequent amendment or addition will be sent to the bidder by e-mail. Prior registration is required for submitting clarification questions on the bidding documents and receiving automatic notice of answers and addenda to the bidding documents. A Complete Set of the Bidding Document can also be downloaded from the link below: (To be provided after NOL)

- Bids must be delivered to the address below on or before 13th December 2026 at 11:00 A.M. Somalia local time. Late Bids will be rejected. The outer Bid envelopes marked “ORIGINAL BID”, and the inner envelopes marked “TECHNICAL PART” will be publicly opened in the presence of the Bidders’ designated representatives and anyone who chooses to attend, at the address below on 13th December 2026 at 11:00 A.M. Somalia local time.

- All Bids must be accompanied by a Bid-Securing Declaration, using the form included in the bidding document

- Attention is drawn to the Procurement Regulations requiring the Borrower to disclose information on the successful bidder’s beneficial ownership, as part of the Contract Award Notice, using the Beneficial Ownership Disclosure Form as included in the bidding document.

- The address (es) referred to above is (are):

| For inspection and clarification of bidding document |  |
| --- | --- |
| Contact person name and title | Procurement Specialist |
| Name of organization | PIU Office (DRIVE Project) Ministry of Finance, Federal Government of Somalia |
| Address | Shangani District, Mogadishu, Somalia |
| Email | bid.drive@piu.mof.gov.so |
| For submission and opening of Technical Parts of the Bids |  |
| Contact person name and title | Procurement Specialist |
| Name of organization | PIU Office (DRIVE Project) Federal Government of Somalia |
| Address | Dahab Apartments, 6th Floor, Next to Ministry of Finance building, Shangani District, Mogadishu, Somalia |
| Date | 13th December 2026 |
| Time | 11:00 A.m. Somalia local time |

Federal Government of Somalia

Ministry of Finance, Somalia

Request for Bids

GOODS

- Supply, Delivery, Installation and Commissioning of Laboratory Equipment for the National Reference Laboratory (NRL) of the Somali Bureau of Standards (SoBS)

| # IFB No: | # SO-MOF-561587-GO-RFB |
| --- | --- |
| # Bid type: | # Single Stage – Two Envelope Bidding with Rated Criteria. |
| # Project: | # De-Risking, Inclusion and Value Enhancement of Pastoral Economies Project (Drive) |
| # Project ID: | # P176517 |
| # Purchaser: | # Ministry of Finance, Somalia |

Table of Contents

PART 1 – Bidding Procedures	3

Section I - Instructions to Bidders	5

Section II - Bid Data Sheet (BDS)	43

Section III - Evaluation and Qualification Criteria	53

Section IV - Bidding Forms	63

Section V - Eligible Countries	87

Section VI - Fraud and Corruption	89

PART 2 – Supply Requirements	91

Section VII - Schedule of Requirements	93

PART 3 – Contract	103

Section VIII - General Conditions of Contract	105

Section IX - Special Conditions of Contract	129

Section X - Contract Forms	141

PART 1 – Bidding Procedures

Section I - Instructions to Bidders

Contents

A.	General	7

1.	Scope of Bid	7

2.	Source of Funds	7

3.	Fraud and Corruption	8

4.	Eligible Bidders	8

5.	Eligible Goods and Related Services	11

B.	Contents of Request for Bids Document	12

6.	Sections of Bidding Document	12

7.	Clarification of the Bidding Document	13

8.	Amendment of Bidding Document	13

C.	Preparation of Bids	14

9.	Cost of Bidding	14

10.	Language of Bid	14

11.	Documents comprising Bid	14

12.	Letters of Bid	16

13.	Alternative Bids	16

14.	Bid prices and Discounts	16

15.	Currencies of Bid and Payment	19

16.	Documents Establishing the Eligibility and Conformity of the Goods and Related Services	19

17.	Documents Establishing the Eligibility and Qualifications of the Bidder	20

18.	Period of Validity of Bids	21

19.	Bid Security	22

20.	Format and Signing of Bid	24

D.	Submission of Bids	24

21.	Sealing and Marking of Bids	24

22.	Deadline for Submission of Bids	25

23.	Late Bids	26

24.	Withdrawal, Substitution, and Modification of Bids	26

E.	Public Opening of Technical Parts of Bids	27

25.	Public Opening of Technical Parts of Bids	27

F.	Evaluation of Bids - General Provisions	29

26.	Confidentiality	29

27.	Clarification of Bids	29

28.	Deviations, Reservations, and Omissions	29

29.	Nonconformities, Errors and Omissions	30

G.	Evaluation of Technical Parts of Bids	30

30.	Evaluation of Technical Parts	30

31.	Determination of Responsiveness	30

32.	Qualification of the Bidders and Detailed Evaluation of the Technical Part	31

H.	Notification of Evaluation of Technical Parts and Public Opening of Financial Parts of Bids	32

33.	Notification of Evaluation of Technical Parts and Public Opening of Financial Parts	32

I.	Evaluation of Financial Parts of Bids	34

34.	Evaluation of Financial Parts	34

35.	Correction of Arithmetic Errors	36

36.	Conversion to Single Currency	36

37.	Margin of  Preference	37

38.	Comparison of Financial Parts	37

39.	Abnormally Low Bids	37

J.	Evaluation of Combined Technical and Financial Parts, Most Advantageous Bid and Notification of Intention to Award	37

40.	Evaluation of combined Technical and Financial Parts	37

41.	Purchaser’s Right to Accept Any Bid, and to Reject Any or All Bids	38

42.	Standstill Period	38

43.	Notification of Intention to Award	38

K.	Award of Contract	39

44.	Award Criteria	39

45.	Purchaser’s Right to Vary Quantities at Time of Award	39

46.	Notification of Award	39

47.	Debriefing by the Purchaser	40

48.	Signing of Contract	41

49.	Performance Security	42

50.	Procurement Related Complaint	42

Section I - Instructions to Bidders

| General |  |
| --- | --- |
| Scope of Bid | - In connection with the Specific Procurement Notice, Request for Bids (RFB), specified in the Bid Data Sheet (BDS), the Purchaser, as specified in the BDS, issues this bidding document for the supply of Goods and, if applicable, any Related Services incidental thereto, as specified in Section VII, Schedule of Requirements. The name, identification and number of lots (contracts) of this RFB are specified in the BDS. - Throughout this bidding document: ### the term “in writing” means communicated in written form (e.g. by mail, e-mail, fax, including, if specified in the BDS, distributed or received through the electronic-procurement system used by the Purchaser), with proof of receipt; ### if the context so requires, “singular” means “plural” and vice versa; and ### “Day” means calendar day, unless otherwise specified as “Business Day”. A Business Day is any day that is an official working day of the Borrower. It excludes the Borrower’s official public holidays. |
| Source of Funds | - The Borrower or Recipient (hereinafter called “Borrower”) specified in the BDS has applied for or received financing (hereinafter called “funds”) from the International Bank for Reconstruction and Development or the International Development Association (hereinafter called “the Bank”) in an amount specified in the BDS, toward the project named in the BDS. The Borrower intends to apply a portion of the funds to eligible payments under the contract for which this bidding document is issued. - Payment by the Bank will be made only at the request of the Borrower and upon approval by the Bank in accordance with the terms and conditions of the Loan (or other financing) Agreement. The Loan (or other financing) Agreement prohibits a withdrawal from the loan account for the purpose of any payment to persons or entities, or for any import of goods, if such payment or import is prohibited by a decision of the United Nations Security Council taken under Chapter VII of the Charter of the United Nations. No party other than the Borrower shall derive any rights from the Loan (or other financing) Agreement or have any claim to the proceeds of the Loan (or other financing). |
| Fraud and Corruption | - The Bank requires compliance with the Bank’s Anti-Corruption Guidelines and its prevailing sanctions policies and procedures as set forth in the WBG’s Sanctions Framework, as set forth in Section VI, Fraud and Corruption. - In further pursuance of this policy, bidders shall permit and shall cause their agents (where declared or not), subcontractors, subconsultants, service providers, suppliers, and personnel, to permit the Bank to inspect all accounts, records and other documents relating to any initial selection process, prequalification process, bid submission, proposal submission, and contract performance (in the case of award), and to have them audited by auditors appointed by the Bank. |
| Eligible Bidders | - A Bidder may be a firm that is a private entity, a state-owned enterprise or institution (subject to ITB 4.6), or any combination of such entities in the form of a joint venture (JV) under an existing agreement or with the intent to enter into such an agreement supported by a letter of intent. In the case of a joint venture, all members shall be jointly and severally liable for the execution of the entire Contract in accordance with the Contract terms. The JV shall nominate a Representative who shall have the authority to conduct all business for and on behalf of any and all the members of the JV during the Bidding process and, in the event the JV is awarded the Contract, during contract execution. Unless specified in the BDS, there is no limit on the number of members in a JV. - A Bidder shall not have a conflict of interest. Any Bidder found to have a conflict of interest shall be disqualified. A Bidder may be considered to have a conflict of interest for the purpose of this Bidding process, if the Bidder: ### directly or indirectly controls, is controlled by or is under common control with another Bidder; or ### receives or has received any direct or indirect subsidy from another Bidder; or ### has the same legal representative as another Bidder; or ### has a relationship with another Bidder, directly or through common third parties, that puts it in a position to influence the Bid of another Bidder, or influence the decisions of the Purchaser regarding this Bidding process; or ### or any of its affiliates participated as a consultant in the preparation of the design or technical specifications of the works that are the subject of the Bid; or ### or any of its affiliates has been hired (or is proposed to be hired) by the Purchaser or Borrower for the Contract implementation; or ### would be providing goods, works, or non-consulting services resulting from or directly related to consulting services for the preparation or implementation of the project specified in the BDS reference ITB 2.1 (the name of the project), that it provided or were provided by any affiliate that directly or indirectly controls, is controlled by, or is under common control with that firm; or ### has a close business or family relationship with a professional staff of the Borrower (or of the project implementing agency, or of a recipient of a part of the loan) who: (i) are directly or indirectly involved in the preparation of the bidding document or specifications of the Contract, and/or the Bid evaluation process of such Contract; or (ii) would be involved in the implementation or supervision of such Contract unless the conflict stemming from such relationship has been resolved in a manner acceptable to the Bank throughout the Bidding process and execution of the Contract. - A firm that is a Bidder (either individually or as a JV member) shall not participate in more than one Bid, except for permitted alternative Bids. This includes participation as a subcontractor. Such participation shall result in the disqualification of all Bids in which the firm is involved. A firm that is not a Bidder or a JV member, may participate as a subcontractor in more than one Bid. - A Bidder may have the nationality of any country, subject to the restrictions pursuant to ITB 4.8. A Bidder shall be deemed to have the nationality of a country if the Bidder is constituted, incorporated or registered in and operates in conformity with the provisions of the laws of that country, as evidenced by its articles of incorporation (or equivalent documents of constitution or association) and its registration documents, as the case may be. This criterion also shall apply to the determination of the nationality of proposed subcontractors or subconsultants for any part of the Contract including related Services. - A Bidder that has been sanctioned by the Bank, pursuant to the Bank’s Anti-Corruption Guidelines, and in accordance with its prevailing sanctions policies and procedures as set forth in the WBG’s Sanctions Framework as described in Section VI paragraph 2.2 d. shall be ineligible to be prequalified for, initially selected for, bid for, propose for, or be awarded a Bank-financed contract or benefit from a Bank-financed contract, financially or otherwise, during such period of time as the Bank shall have determined. The list of debarred firms and individuals is available at the electronic address specified in the BDS. - Bidders that are state-owned enterprises or institutions in the Purchaser’s Country may be eligible to compete and be awarded a Contract(s) only if they can establish, in a manner acceptable to the Bank, that they (i) are legally and financially autonomous (ii) operate under commercial law, and (iii) are not under supervision of the Purchaser. - A Bidder shall not be under suspension from Bidding by the Purchaser as the result of the operation of a Bid–Securing Declaration or Proposal-Securing Declaration. - Firms and individuals may be ineligible if so indicated in Section V, Eligible Countries, and: - as a matter of law or official regulations, the Borrower’s country prohibits commercial relations with that country, provided that the Bank is satisfied that such exclusion does not preclude effective competition for the supply of goods or the contracting of works or services required; or - by an act of compliance with a decision of the United Nations Security Council taken under Chapter VII of the Charter of the United Nations, the Borrower’s country prohibits any import of goods or contracting of works or services from that country, or any payments to any country, person, or entity in that country. - A Bidder shall provide such documentary evidence of eligibility satisfactory to the Purchaser, as the Purchaser shall reasonably request. - A firm that is under a sanction of debarment by the Borrower from being awarded a contract is eligible to participate in this procurement, unless the Bank, at the Borrower’s request, is satisfied that the debarment; - relates to fraud or corruption; and - followed a judicial or administrative proceeding that afforded the firm adequate due process. |
| Eligible Goods and Related Services | - All the Goods and Related Services to be supplied under the Contract and financed by the Bank may have their origin in any country in accordance with Section V, Eligible Countries. - For purposes of this ITB, the term “goods” includes commodities, raw material, machinery, equipment, and industrial plants; and “related services” includes services such as insurance, installation, training, and initial maintenance. - The term “origin” means the country where the goods have been mined, grown, cultivated, produced, manufactured or processed; or, through manufacture, processing, or assembly, another commercially recognized article results that differs substantially in its basic characteristics from its components. |
| Contents of Request for Bids Document |  |
| Sections of Bidding Document | - The bidding document consist of Parts 1, 2, and 3, which include all the sections indicated below, and should be read in conjunction with any addenda issued in accordance with ITB 8. PART 1 Bidding Procedures - Section I - Instructions to Bidders (ITB) - Section II - Bidding Data Sheet (BDS) - Section III - Evaluation and Qualification Criteria - Section IV - Bidding Forms - Section V - Eligible Countries - Section VI - Fraud and Corruption |
|  | PART 2 Supply Requirements - Section VII - Schedule of Requirements PART 3 Contract - Section VIII - General Conditions of Contract - Section IX - Special Conditions of Contract - Section X - Contract Forms |
|  | - The Specific Procurement Notice - Request for Bids (RFB) issued by the Purchaser is not part of this bidding document. - Unless obtained directly from the Purchaser, the Purchaser is not responsible for the completeness of the document, responses to requests for clarification, the Minutes of the pre-Bid meeting (if any), or addenda to the bidding document in accordance with ITB 8. In case of any contradiction, documents obtained directly from the Purchaser shall prevail. - The Bidder is expected to examine all instructions, forms, terms, and specifications in the bidding document and to furnish with its Bid all information or documentation as is required by the bidding document. |
| Clarification of the Bidding Document | - A Bidder requiring any clarification of the bidding document shall contact the Purchaser in writing at the Purchaser’s address specified in the BDS. The Purchaser will respond in writing to any request for clarification, provided that such request is received prior to the deadline for submission of Bids within a period specified in the BDS. The Purchaser shall forward copies of its response to all Bidders who have acquired the bidding document in accordance with ITB 6.3, including a description of the inquiry but without identifying its source. If so specified in the BDS, the Purchaser shall also promptly publish its response at the web page identified in the BDS. Should the clarification result in changes to the essential elements of the bidding document, the Purchaser shall amend the bidding document following the procedure under ITB 8 and ITB 22.2. |
| Amendment of Bidding Document | - At any time prior to the deadline for submission of Bids, the Purchaser may amend the bidding document by issuing addenda. - Any addendum issued shall be part of the bidding document and shall be communicated in writing to all who have obtained the bidding document from the Purchaser in accordance with ITB 6.3. The Purchaser shall also promptly publish the addendum on the Purchaser’s web page in accordance with ITB 7.1. - To give prospective Bidders reasonable time in which to take an addendum into account in preparing their Bids, the Purchaser may, at its discretion, extend the deadline for the submission of Bids, pursuant to ITB 22.2. |
| Preparation of Bids |  |
| Cost of Bidding | - The Bidder shall bear all costs associated with the preparation and submission of its Bid, and the Purchaser shall not be responsible or liable for those costs, regardless of the conduct or outcome of the Bidding process. |
| Language of Bid | - The Bid, as well as all correspondence and documents relating to the Bid exchanged by the Bidder and the Purchaser, shall be written in the language specified in the BDS. Supporting documents and printed literature that are part of the Bid may be in another language provided they are accompanied by an accurate translation of the relevant passages into the language specified in the BDS, in which case, for purposes of interpretation of the Bid, such translation shall govern. |
| Documents comprising Bid | - The Bid shall comprise two Parts, namely the Technical Part and the Financial Part. These two Parts shall be submitted simultaneously in two separate sealed envelopes (two-envelope Bidding process). One envelope shall contain only information relating to the Technical Part and the other, only information relating to the Financial Part. These two envelopes shall be enclosed in a separate sealed outer envelope marked “Original Bid”. - The Technical Part shall contain the following: ### Letter of Bid - Technical Part: prepared in accordance with ITB 12; ### Bid Security or Bid-Securing Declaration: in accordance with ITB 19.1; ### Alternative Bid - Technical Part: if permissible in accordance with ITB 13, the Technical Part of any Alternative Bid; ### Authorization: written confirmation authorizing the signatory of the Bid to commit the Bidder, in accordance with ITB 20.3; ### Bidder’s Eligibility: documentary evidence in accordance with ITB 17 establishing the Bidder’s eligibility to Bid; ### Qualifications: documentary evidence in accordance with ITB 17 establishing the Bidder’s qualifications to perform the Contract if its Bid is accepted; ### Eligibility of Goods and Related Services: documentary evidence in accordance with ITB 16, establishing the eligibility of the Goods and Related Services to be supplied by the Bidder; ### Conformity: documentary evidence in accordance with ITB 16, that the Goods and Related Services conform to the bidding document; ### any other document required in the BDS. - The Financial Part envelope shall contain the following: - Letter of Bid – Financial Part: prepared in accordance with ITB 12 and ITB 14; - Price Schedules: completed prepared in accordance with ITB 12 and ITB 14; ### Alternative Bid - Financial Part; if permissible in accordance with ITB 13, the Financial Part of any Alternative Bid; ### any other document required in the BDS. - The Technical Part shall not include any financial information related to the Bid price. Where material financial information related to the Bid price is contained in the Technical Part the Bid shall be declared non-responsive. - In addition to the requirements under ITB 11.2, Bids submitted by a JV shall include a copy of the Joint Venture Agreement entered into by all members. Alternatively, a letter of intent to execute a Joint Venture Agreement in the event of a successful Bid shall be signed by all members and submitted with the Bid, together with a copy of the proposed Agreement. - The Bidder shall furnish in the Letter of Bid information on commissions and gratuities, if any, paid or to be paid to agents or any other party relating to this Bid. |
| Letters of Bid | - The Bidder shall prepare the Letter of Bid – Technical Part, and Letter of Bid – Financial Part using the relevant forms furnished in Section IV, Bidding Forms. The forms must be completed without any alterations to the text, and no substitutes shall be accepted except as provided under ITB 20.3. All blank spaces shall be filled in with the information requested. |
| Alternative Bids | - Unless otherwise specified in the BDS, Alternative Bids shall not be considered. |
| Bid prices and Discounts | - The prices and discounts quoted by the Bidder in the Letter of Bid - Financial Part and in the Price Schedules shall conform to the requirements specified below. - All lots (contracts) and items must be listed and priced separately in the Price Schedules. - The price to be quoted in the Letter of Bid - Financial Part, in accordance with ITB 12.1 shall be the total price of the Bid, excluding any discounts offered. - The Bidder shall quote any discounts and indicate the methodology for their application in the Letter of Bid - Financial Part, in accordance with ITB 12.1. - Prices quoted by the Bidder shall be fixed during the Bidder’s performance of the Contract and not subject to variation on any account, unless otherwise specified in the BDS. A Bid submitted with an adjustable price quotation shall be treated as nonresponsive and shall be rejected, pursuant to ITB 31. However, if in accordance with the BDS, prices quoted by the Bidder shall be subject to adjustment during the performance of the Contract, a Bid submitted with a fixed price quotation shall not be rejected, but the price adjustment shall be treated as zero. - If so specified in ITB 1.1, Bids are being invited for individual lots (contracts) or for any combination of lots (packages). Unless otherwise specified in the BDS, prices quoted shall correspond to 100% of the items specified for each lot and to 100% of the quantities specified for each item of a lot. Bidders wishing to offer discounts for the award of more than one Contract shall specify in their Bid the price reductions applicable to each package, or alternatively, to individual Contracts within the package. However, discounts that are conditional on the award of more that one lot will not be considered for bid evaluation purpose. - The terms EXW, CIP, and other similar terms shall be governed by the rules prescribed in the current edition of Incoterms, published by the International Chamber of Commerce, as specified in the BDS. - Prices shall be quoted as specified in each Price Schedule included in Section IV, Bidding Forms. The disaggregation of price components is required solely for the purpose of facilitating the comparison of Bids by the Purchaser. This shall not in any way limit the Purchaser’s right to contract on any of the terms offered. In quoting prices, the Bidder shall be free to use transportation through carriers registered in any eligible country, in accordance with Section V, Eligible Countries. Similarly, the Bidder may obtain insurance services from any eligible country in accordance with Section V, Eligible Countries. Prices shall be entered in the following manner: ### For Goods manufactured in the Purchaser’s Country: - the price of the Goods quoted EXW (ex-works, ex-factory, ex warehouse, ex showroom, or off-the-shelf, as applicable), including all customs duties and sales and other taxes already paid or payable on the components and raw material used in the manufacture or assembly of the Goods; - any Purchaser’s Country sales tax and other taxes which will be payable on the Goods if the Contract is awarded to the Bidder; and - the price for inland transportation, insurance, and other local services required to convey the Goods to their final destination (Project Site) specified in the BDS. ### For Goods manufactured outside the Purchaser’s Country, to be imported: - the price of the Goods, quoted CIP named place of destination, in the Purchaser’s Country, as specified in the BDS; - the price for inland transportation, insurance, and other local services required to convey the Goods from the named place of destination to their final destination (Project Site) specified in the BDS; ### For Goods manufactured outside the Purchaser’s Country, already imported: - the price of the Goods, including the original import value of the Goods; plus any mark-up (or rebate); plus any other related local cost, and custom duties and other import taxes already paid or to be paid on the Goods already imported. - the custom duties and other import taxes already paid (need to be supported with documentary evidence) or to be paid on the Goods already imported; - the price of the Goods, obtained as the difference between (i) and (ii) above; - any Purchaser’s Country sales and other taxes which will be payable on the Goods if the Contract is awarded to the Bidder; and - the price for inland transportation, insurance, and other local services required to convey the Goods to their final destination (Project Site) specified in the BDS. ### for Related Services, other than inland transportation and other services required to convey the Goods to their final destination, whenever such Related Services are specified in the Schedule of Requirements: - the price of each item comprising the Related Services (inclusive of any applicable taxes). |
| Currencies of Bid and Payment | - The currency(ies) of the Bid and the currency(ies) of payments shall be the same. The Bidder shall quote in the currency of the Purchaser’s Country the portion of the Bid price that corresponds to expenditures incurred in the currency of the Purchaser’s country, unless otherwise specified in the BDS. - The Bidder may express the Bid price in any currency. If the Bidder wishes to be paid in a combination of amounts in different currencies, it may quote its price accordingly but shall use no more than three foreign currencies in addition to the currency of the Purchaser’s Country. |
| Documents Establishing the Eligibility and Conformity of the Goods and Related Services | - To establish the eligibility of the Goods and Related Services in accordance with ITB 5, Bidders shall complete the country of origin declarations in the Price Schedule Forms, included in Section IV, Bidding Forms. - To establish the conformity of the Goods and Related Services to the bidding document, the Bidder shall furnish as part of its Bid the documentary evidence that the Goods conform to the technical specifications and standards specified in Section VII, Schedule of Requirements. - The documentary evidence may be in the form of literature, drawings or data, and shall consist of a detailed item by item description of the essential technical and performance characteristics of the Goods and Related Services, demonstrating substantial responsiveness of the Goods and Related Services to the technical specification, and if applicable, a statement of deviations and exceptions to the provisions of the Section VII, Schedule of Requirements. - The Bidder shall also furnish a list giving full particulars, including available sources and current prices of spare parts, special tools, etc., necessary for the proper and continuing functioning of the Goods during the period specified in the BDS following commencement of the use of the goods by the Purchaser. - Standards for workmanship, process, material, and equipment, as well as references to brand names or catalogue numbers specified by the Purchaser in the Schedule of Requirements, are intended to be descriptive only and not restrictive. The Bidder may offer other standards of quality, brand names, and/or catalogue numbers, provided that it demonstrates, to the Purchaser’s satisfaction, that the substitutions ensure substantial equivalence or are superior to those specified in the Section VII, Schedule of Requirements. |
| Documents Establishing the Eligibility and Qualifications of the Bidder | - To establish Bidder’s eligibility in accordance with ITB 4, Bidders shall complete the Letter of Bid – Technical Part, included in Section IV, Bidding Forms. - The documentary evidence of the Bidder’s qualifications to perform the Contract, if its Bid is accepted, shall establish to the Purchaser’s satisfaction: - that, if required in the BDS, a Bidder that does not manufacture or produce the Goods it offers to supply shall submit the Manufacturer’s Authorization using the form included in Section IV, Bidding Forms to demonstrate that it has been duly authorized by the manufacturer or producer of the Goods to supply these Goods in the Purchaser’s Country; - that, if required in the BDS, in case of a Bidder not doing business within the Purchaser’s Country, the Bidder is or will be (if awarded the Contract) represented by an Agent in the country equipped and able to carry out the Supplier’s maintenance, repair and spare parts-stocking obligations prescribed in the Conditions of Contract and/or Technical Specifications; and - that the Bidder meets each of the qualification criterion specified in Section III, Evaluation and Qualification Criteria. |
| Period of Validity of Bids | - Bids shall remain valid until the date specified in the BDS or any extended date if amended by the Purchaser in accordance with ITB 8. A Bid that is not valid until the date specified in the BDS, or any extended date if amended by the Purchaser in accordance with ITB 8, shall be rejected by the Purchaser as nonresponsive. - In exceptional circumstances, prior to the expiry of the Bid validity, the Purchaser may request Bidders to extend the period of validity of their Bids. The request and the responses shall be made in writing. If a Bid Security is requested (in accordance with ITB 19), it shall also be extended for a corresponding period. A Bidder may refuse the request without forfeiting its Bid Security. A Bidder granting the request shall not be required or permitted to modify its Bid, except as provided in ITB 18.3. - If the award is delayed by a period exceeding fifty-six (56) days beyond the expiry of the initial Bid validity, the Contract price shall be determined as follows: - In the case of fixed price contracts, the Contract price shall be the Bid price adjusted by the factor specified in the BDS. - In the case of adjustable price contracts, no adjustment shall be made. - In any case, Bid evaluation shall be based on the Bid price without taking into consideration the applicable correction from those indicated above. |
| Bid Security | - The Bidder shall furnish, as part of the Technical Part of its Bid, either a Bid-Securing Declaration or a Bid Security, as specified in the BDS, in original form and, in the case of a Bid security, in the amount and currency specified in the BDS. - A Bid Securing Declaration shall use the form included in Section IV, Bidding Forms. - If a Bid Security is specified pursuant to ITB 19.1, the Bid security shall be a demand guarantee in any of the following forms at the Bidder’s option: ### an unconditional guarantee issued by a bank or non-bank financial institution (such as an insurance, bonding or surety company); ### an irrevocable letter of credit; ### a cashier’s or certified check; or ### another security specified in the BDS, from a reputable source from an eligible country. If an unconditional guarantee is issued by a non-bank financial institution located outside the Purchaser’s Country the issuing non-bank financial institution shall have a correspondent financial institution located in the Purchaser’s Country to make it enforceable unless the Purchaser has agreed in writing, prior to Bid submission, that a correspondent financial institution is not required. In the case of a bank guarantee, the Bid security shall be submitted either using the Bid Security Form included in Section IV, Bidding Forms, or in another substantially similar format approved by the Purchaser prior to Bid submission. The Bid security shall be valid for twenty-eight (28) days beyond the original date of expiry of the Bid validity, or beyond any extended date if requested under ITB 18.2. - If a Bid Security is specified pursuant to ITB 19.1, any Bid not accompanied by a substantially responsive Bid Security shall be rejected by the Purchaser as non-responsive. - If a Bid Security is specified pursuant to ITB 19.1, the Bid Security of unsuccessful Bidders shall be returned as promptly as possible upon the successful Bidder’s signing the contract and furnishing the Performance Security pursuant to ITB 49. - The Bid Security of the successful Bidder shall be returned as promptly as possible once the successful Bidder has signed the Contract and furnished the required performance security. - The Bid Security may be forfeited: ### if a Bidder withdraws its Bid prior to the expiry date of Bid validity specified by the Bidder on the Letter of Bid or any extended date provided by the Bidder ; or ### if the successful Bidder fails to: ### sign the Contract in accordance with ITB 48; or ### furnish a performance security in accordance with ITB 49. - The Bid Security or Bid-Securing Declaration of a JV must be in the name of the JV that submits the Bid. If the JV has not been legally constituted into a legally enforceable JV at the time of Bidding, the Bid security or Bid-Securing Declaration shall be in the names of all future members as named in the letter of intent referred to in ITB 4.1 and ITB 11.5. - If a Bid security is not required in the BDS, pursuant to ITB 19.1, and - if a Bidder withdraws its Bid during the period of Bid validity specified by the Bidder on the Letter of Bid, or any extended date provided by the Bidder, or - if the successful Bidder fails to: sign the Contract in accordance with ITB 48; or furnish a performance security in accordance with ITB 49; the Borrower may, if provided for in the BDS, declare the Bidder ineligible to be awarded a contract by the Purchaser for a period of time as stated in the BDS. |
| Format and Signing of Bid | - The Bidder shall prepare the Bid, in accordance with ITB 11 and ITB 21. - Bidders shall mark as “Confidential” information in their Bids which is confidential to their business. This may include proprietary information, trade secrets, or commercial or financially sensitive information. - The original and all copies of the Bid shall be typed or written in indelible ink and shall be signed by a person duly authorized to sign on behalf of the Bidder. This authorization shall consist of a written confirmation as specified in the BDS and shall be attached to the Bid. The name and position held by each person signing the authorization must be typed or printed below the signature. All pages of the Bid where entries or amendments have been made shall be signed or initialed by the person signing the Bid. - In case the Bidder is a JV, the Bid shall be signed by an authorized representative of the JV on behalf of the JV, and so as to be legally binding on all the members as evidenced by a power of attorney signed by their legally authorized representatives. - Any inter-lineation, erasures, or overwriting shall be valid only if they are signed or initialed by the person signing the Bid. |
| Submission of Bids |  |
| Sealing and Marking of Bids | - The Bidder shall deliver the Bid in two separate, sealed envelopes (the Technical Part and the Financial Part). These two envelopes shall be enclosed in a sealed outer envelope marked “Original Bid”. - In addition, the Bidder shall submit copies of the Bid in the number specified in the BDS. Copies of the Technical Part shall be placed in a separate sealed envelope marked “Copies: Technical Part”. Copies of the Financial Part shall be placed in a separate sealed envelope marked “Copies: Financial Part”. The Bidder shall place both of these envelopes in a separate, sealed outer envelope marked “Bid Copies”. In the event of any discrepancy between the original and the copies, the original shall prevail. If alternative Bids are permitted in accordance with ITB 13, the alternative Bids shall be submitted as follows: the original of the alternative Bid Technical Part shall be placed in a sealed envelope marked “Alternative Bid – Technical Part” and the Financial Part shall be placed in a sealed envelope marked “Alternative Bid – Financial Part” and these two separate sealed envelopes then enclosed within a sealed outer envelope marked “Alternative Bid – Original”, the copies of the alternative Bid will be placed in separate sealed envelopes marked “Alternative Bid – Copies Of Technical Part”, and “Alternative Bid – Copies Of Financial Part” and enclosed in a separate sealed outer envelope marked “Alternative Bid - Copies”. - The envelopes marked “Original Bid” and “Bid Copies” (and, if appropriate, a third envelope marked “Alternative Bid”) shall be enclosed in a separate sealed outer envelope for submission to the Purchaser. - All inner and outer envelopes, shall: - bear the name and address of the Bidder; - be addressed to the Purchaser in accordance with ITB 22.1; - bear the specific identification of this Bidding process indicated in ITB 1.1; and - bear a warning not to open before the time and date for Bid opening. - If all envelopes are not sealed and marked as required, the Purchaser will assume no responsibility for the misplacement or premature opening of the Bid. |
| Deadline for Submission of Bids | - Bids must be received by the Purchaser at the address and no later than the date and time specified in the BDS. When so specified in the BDS, Bidders shall have the option of submitting their Bids electronically. Bidders submitting Bids electronically shall follow the electronic Bid submission procedures specified in the BDS. - The Purchaser may, at its discretion, extend the deadline for the submission of Bids by amending the bidding document in accordance with ITB 8, in which case all rights and obligations of the Purchaser and Bidders previously subject to the deadline shall thereafter be subject to the deadline as extended. |
| Late Bids | - The Purchaser shall not consider any Bid that arrives after the deadline for submission of Bids, in accordance with ITB 22. Any Bid received by the Purchaser after the deadline for submission of Bids shall be declared late, rejected, and returned unopened to the Bidder. |
| Withdrawal, Substitution, and Modification of Bids | - A Bidder may withdraw, substitute, or modify its Bid after it has been submitted by sending a written notice, duly signed by an authorized representative, and shall include a copy of the authorization (the power of attorney) in accordance with ITB 20.3, (except that withdrawal notices do not require copies ). The corresponding substitution or modification of the Bid must accompany the respective written notice. All notices must be: - prepared and submitted in accordance with ITB 20 and ITB 21 (except that withdrawal notices do not require copies), and in addition, the respective envelopes shall be clearly marked “Withdrawal,” “Substitution,” or “Modification;” and - received by the Purchaser prior to the deadline prescribed for submission of Bids, in accordance with ITB 22. - Bids requested to be withdrawn in accordance with ITB 24.1 shall be returned unopened to the Bidders. - No Bid may be withdrawn, substituted, or modified in the interval between the deadline for submission of Bids and the expiration of the period of Bid validity specified by the Bidder on the Letter of Bid -Technical Part and repeated in the Letter of Bid - Financial Part, or any extension thereof. |
| Public Opening of Technical Parts of Bids |  |
| Public Opening of Technical Parts of Bids | - Except as in the cases specified in ITB 23 and ITB 24.2, the Purchaser shall, at this Bid opening, publicly open and read out, in accordance with this ITB, all bids received by the deadline at the date, time and place specified in the BDS in the presence of Bidders’ designated representatives and anyone who chooses to attend. Any specific electronic Bid opening procedures required if electronic Bidding is permitted in accordance with ITB 22.1, shall be as specified in the BDS. - First, the written notice of withdrawal in the envelopes marked “Withdrawal” shall be opened and read out and the envelope with the corresponding Bid shall not be opened, but returned to the Bidder. If the withdrawal envelope does not contain a copy of the “power of attorney” confirming the signature as a person duly authorized to sign on behalf of the Bidder, the corresponding Bid will be opened. No Bid withdrawal shall be permitted unless the corresponding withdrawal notice contains a valid authorization to request the withdrawal and is read out at Bid opening. - Next, envelopes marked “Substitution” shall be opened and read out and exchanged with the corresponding Bid being substituted, and the substituted Bid shall not be opened, but returned to the Bidder. No Bid substitution shall be permitted unless the corresponding substitution notice contains a valid authorization to request the substitution and is read out at Bid opening. - Next, envelopes marked “Modification” shall be opened and read out with the corresponding Bid. No Bid modification shall be permitted unless the corresponding modification notice contains a valid authorization to request the modification and is read out at Bid opening. Only Bids that are opened and read out at Bid opening shall be considered further. - Next, all other envelopes marked “Technical Part” shall be opened one at a time. All envelopes marked “Financial Part” shall remain sealed, and kept by the Purchaser in safe custody until they are opened, at a later public opening, following the evaluation of the Technical Part of the Bids. On opening the envelopes marked “Technical Part” the Purchaser shall read out: the name of the Bidder and whether there is a modification; and Alternative Bid the presence or absence of a Bid Security, if required and any other details as the Purchaser may consider appropriate. - Only Technical Parts of Bids and Alternative Bid - Technical Parts that are read out at Bid opening shall be considered further in the evaluation. The Letter of Bid – Technical Part and the separate sealed envelope marked “Financial Part” are to be initialed by representatives of the Purchaser attending Bid opening in the manner specified in the BDS. - At the Bid opening the Purchaser shall neither discuss the merits of any Bid nor reject any Bid (except for late Bids, in accordance with ITB 23.1). - Following the opening of the Technical Parts of the Bid the Purchaser shall prepare a record that shall include, as a minimum: - the name of the Bidder and whether there is a withdrawal, substitution, or modification; - the presence or absence of a duly sealed envelope marked “Financial Part”; - the presence or absence of a Bid Security or Bid-Securing Declaration; and - if applicable, any Alternative Bid - Technical Part; - The Bidders’ representatives who are present shall be requested to sign the record. The omission of a Bidder’s signature on the record shall not invalidate the contents and effect of the record. A copy of the record shall be distributed to all Bidders. |
| Evaluation of Bids - General Provisions |  |
| Confidentiality | - Information relating to the evaluation of the Technical Part shall not be disclosed to Bidders or any other persons not officially concerned with the Bidding process until the notification of evaluation of the Technical Part in accordance with ITB 33. Information relating to the evaluation of Financial Part, the evaluation of combined Technical Part and Financial Part, and recommendation of contract award shall not be disclosed to Bidders or any other persons not officially concerned with the RFB process until the Notification of Intention to Award the Contract is transmitted to Bidders in accordance with ITB 43. - Any effort by a Bidder to influence the Purchaser in the evaluation or contract award decisions may result in the rejection of its Bid. - Notwithstanding ITB 26.2, from the time of Bid opening to the time of Contract Award, if any Bidder wishes to contact the Purchaser on any matter related to the Bidding process, it should do so in writing. |
| Clarification of Bids | - To assist in the examination, evaluation, comparison of the Bids, and qualification of the Bidders, the Purchaser may, at its discretion, ask any Bidder for a clarification of its Bid. Any clarification submitted by a Bidder in respect to its Bid and that is not in response to a request by the Purchaser shall not be considered. The Purchaser’s request for clarification and the response shall be in writing. No change, including any voluntary increase or decrease, in the prices or substance of the Bid shall be sought, offered, or permitted, except to confirm the correction of arithmetic errors discovered by the Purchaser in the Evaluation of the Bids, in accordance with ITB 35. - If a Bidder does not provide clarifications of its Bid by the date and time set in the Purchaser’s request for clarification, its Bid may be rejected. |
| Deviations, Reservations, and Omissions | - During the evaluation of Bids, the following definitions apply: - “Deviation” is a departure from the requirements specified in the bidding document; - “Reservation” is the setting of limiting conditions or withholding from complete acceptance of the requirements specified in the bidding document; and - “Omission” is the failure to submit part or all of the information or documentation required in the bidding document. |
| Nonconformities, Errors and Omissions | - Provided that a Bid is substantially responsive, the Purchaser may waive any nonconformities in the Bid. - Provided that a Bid is substantially responsive, the Purchaser may request that the Bidder submit the necessary information or documentation, within a reasonable period of time, to rectify nonmaterial nonconformities or omissions in the Bid related to documentation requirements. Such omission shall not be related to any aspect of the price of the Bid. Failure of the Bidder to comply with the request may result in the rejection of its Bid. |
| Evaluation of Technical Parts of Bids |  |
| Evaluation of Technical Parts | - In evaluating the Technical Parts of each Bid, the Purchaser shall use the criteria and methodologies listed in ITB 31, ITB 32, the BDS, if applicable, and Section III, Evaluation and Qualification Criteria. No other evaluation criteria or methodologies shall be permitted. |
| Determination of Responsiveness | - The Purchaser’s determination of a Bid’s responsiveness is to be based on the contents of the Bid itself, as defined in ITB 11. A substantially responsive Bid is one that meets the requirements of the bidding document without material deviation, reservation, or omission. A material deviation, reservation, or omission is one that: ### if accepted, would: ### affect in any substantial way the scope, quality, or performance of the Goods and Related Services specified in the Contract; or ### limit in any substantial way, inconsistent with the bidding document, the Purchaser’s rights or the Bidder’s obligations under the Contract; or ### if rectified, would unfairly affect the competitive position of other Bidders presenting substantially responsive Bids. - The Purchaser shall examine the technical aspects of the Bid submitted in accordance with ITB 16 and ITB 17, in particular, to confirm that all requirements of Section VII, Schedule of Requirements have been met without any material deviation or reservation, or omission. - If a Bid is not substantially responsive to the requirements of bidding document, it shall be rejected by the Purchaser and may not subsequently be made responsive by correction of the material deviation, reservation, or omission. |
| Qualification of the Bidders and Detailed Evaluation of the Technical Part | - The Purchaser shall determine, to its satisfaction, whether all eligible Bidders, whose Bids have been determined to be substantially responsive to the bidding document, meet the Qualification Criteria specified in Section III, Evaluation and Qualification Criteria. - The determination shall be based upon an examination of the documentary evidence of the Bidder’s qualifications submitted by the Bidder, pursuant to ITB 17. The determination shall not take into consideration the qualifications of other firms such as the Bidder’s subsidiaries, parent entities, affiliates, subcontractors (other than specialized subcontractors if permitted in the bidding document), or any other firm different from the Bidder. - Prior to Contract award, the Purchaser will verify that the successful Bidder (including each member of a JV) is not disqualified by the Bank due to noncompliance with contractual SEA/SH prevention and response obligations. The Purchaser will conduct the same verification for each subcontractor proposed by the successful Bidder. If any proposed subcontractor does not meet the requirement, the Purchaser will require the Bidder to propose a replacement subcontractor. - Only substantially responsive bids submitted by eligible and qualified bidders shall proceed to the detailed technical evaluation to assess adequacy of the Technical Part followed by evaluation applying technical factors/subfactors and corresponding scores as specified in the BDS. |
| Notification of Evaluation of Technical Parts and Public Opening of Financial Parts of Bids |  |
| Notification of Evaluation of Technical Parts and Public Opening of Financial Parts | - Following the completion of the evaluation of the Technical Parts of the Bids, and the Bank has issued its no objection (if applicable), the Purchaser shall notify in writing those Bidders who have failed to meet the Qualification Criteria and/or whose Bids were considered non-responsive to the requirements in the bidding document, advising them of the following information: ### the grounds on which their Technical Part of Bid failed to meet the requirements of the bidding document; ### their envelope marked “Financial Part” will be returned to them unopened after the completion of the bid evaluation process and the signing of the Contract; ### notify them of the date, time and location of the public opening of the envelopes marked ‘Financial Part”. - The Purchaser shall, simultaneously, notify in writing those Bidders whose Technical Parts have been evaluated as substantially responsive to the bidding document and met the Qualification Criteria, advising them of the following information: ### their Bid has been evaluated as substantially responsive to the bidding document and met the Qualification Criteria; and ### their envelope marked “Financial Part” will be opened at the public opening of Financial Parts; ### notify them of the date, time and location of the public opening of the envelopes marked “Financial Part”. - The opening date shall be not less than ten (10) Business Days from the date of notification of the results of the technical evaluation, specified in ITB 33.1 and 33.2. However, if the Purchaser receives a complaint on the results of the technical evaluation within the ten (10) Business Days, the opening date shall be subject to ITB 50.1. The Financial Part of the Bid shall be opened publicly in the presence of Bidders’ designated representatives and anyone who chooses to attend. - At this public opening the Financial Parts will be opened by the Purchaser in the presence of Bidders, or their designated representatives and anyone else who chooses to attend. Bidders who met the Qualification Criteria and whose Bids were evaluated as substantially responsive will have their envelopes marked “Financial Part” opened at the second public opening. Each of these envelopes marked “Financial Part” shall be inspected to confirm that they have remained sealed and unopened. These envelopes shall then be opened by the Purchaser. The Purchaser shall read out the names of each Bidder, the technical score and the total Bid prices, per lot (contract) if applicable, including any discounts and Alternative Bid - Financial Part, and any other details as the Purchaser may consider appropriate. - Only envelopes of Financial Part of Bids, Financial Parts of Alternative Bids and discounts that are opened and read out at Bid opening shall be considered further for evaluation. The Letter of Bid - Financial Part and the Price Schedules are to be initialed by a representative of the Purchaser attending the Bid opening in the manner specified in the BDS. - The Purchaser shall neither discuss the merits of any Bid nor reject any envelopes marked “Financial Part”. - The Purchaser shall prepare a record of the Financial Part of the Bid opening that shall include, as a minimum: (a) the name of the Bidder whose Financial Part was opened; (b) the Bid price, per lot (contract) if applicable, including any discounts, (c) if applicable, any Alternative Bid - Financial Part. - The Bidders whose envelopes marked ‘Financial Part” have been opened or their representatives who are present shall be requested to sign the record. The omission of a Bidder’s signature on the record shall not invalidate the contents and effect of the record. A copy of the record shall be distributed to all Bidders. |
| Evaluation of Financial Parts of Bids |  |
| Evaluation of Financial Parts | - Provided that a Bid is substantially responsive, the Purchaser shall rectify quantifiable nonmaterial nonconformities related to the Bid Price. To this effect, the Bid Price shall be adjusted, for comparison purposes only, to reflect the price of a missing or non-conforming item or component, by adding the average price of the item or component quoted by substantially responsive Bidders. If the price of the item or component cannot be derived from the price of other substantially responsive Bids, the Purchaser shall use its best estimate. - To evaluate the Financial Part of each Bid, the Purchaser shall consider the following: ### evaluation will be done for Items or Lots (contracts), as specified in the BDS; and the Bid Price as quoted in accordance with ITB 14; ### price adjustment for correction of arithmetic errors in accordance with ITB 35.1; ### price adjustment due to discounts offered in accordance with ITB 14.4; ### converting the amount resulting from applying (a) to (c) above, if relevant, to a single currency in accordance with ITB 36; ### price adjustment due to quantifiable nonmaterial nonconformities in accordance with ITB 34.1; and ### the additional evaluation factors specified in Section III, Evaluation and Qualification Criteria. - The estimated effect of the price adjustment provisions of the Conditions of Contract, applied over the period of execution of the Contract, shall not be taken into account in Bid evaluation. - If this bidding document allows Bidders to quote separate prices for different lots (contracts), each lot will be evaluated separately to determine the Most Advantageous Bid using the methodology specified in Section III, Evaluation and Qualification Criteria. Discounts that are conditional on the award of more than one lot or slice shall not be considered for Bid evaluation. - The Purchaser’s evaluation of a Bid will exclude and not take into account: ### in the case of Goods manufactured in the Purchaser’s Country, sales and other similar taxes, which will be payable on the goods if a contract is awarded to the Bidder; ### in the case of Goods manufactured outside the Purchaser’s Country, already imported or to be imported, customs duties and other import taxes levied on the imported Good, sales and other similar taxes, which will be payable on the Goods if the contract is awarded to the Bidder; ### any allowance for price adjustment during the period of execution of the contract, if provided in the Bid. - The Purchaser’s evaluation of a Bid may require the consideration of other factors, in addition to the Bid price quoted in accordance with ITB 14. These factors may be related to the characteristics, performance, and terms and conditions of purchase of the Goods and Related Services. The effect of the factors selected, if any, shall be expressed in monetary terms to facilitate comparison of Bids, unless otherwise specified in the BDS from amongst those set out in Section III, Evaluation and Qualification Criteria. The criteria and methodologies to be used shall be as specified in ITB 34.2 (f). |
| Correction of Arithmetic Errors | - In evaluating the Financial Part of each Bid, the Purchaser shall correct arithmetic errors on the following basis: ### if there is a discrepancy between the unit price and the line item total that is obtained by multiplying the unit price by the quantity, the unit price shall prevail and the line item total shall be corrected, unless in the opinion of the Purchaser there is an obvious misplacement of the decimal point in the unit price, in which case the line item total as quoted shall govern and the unit price shall be corrected; ### if there is an error in a total corresponding to the addition or subtraction of subtotals, the subtotals shall prevail and the total shall be corrected; and ### if there is a discrepancy between words and figures, the amount in words shall prevail, unless the amount expressed in words is related to an arithmetic error, in which case the amount in figures shall prevail subject to (a) and (b) above. - Bidders shall be requested to accept correction of arithmetic errors. Failure to accept the correction in accordance with ITB 35.1, shall result in the rejection of the Bid. |
| Conversion to Single Currency | - For evaluation and comparison purposes, the currency(ies) of the Bids shall be converted in a single currency as specified in the BDS. |
| Margin of Preference | - Unless otherwise specified in the BDS, a margin of preference shall not apply. |
| Comparison of Financial Parts | - The Purchaser shall compare the evaluated costs of the Bids to determine the Bid that has the lowest evaluated cost. The comparison shall be on the basis of CIP (place of final destination) prices for imported goods and EXW prices, plus cost of inland transportation and insurance to place of destination, for goods manufactured within the Borrower’s country, together with prices for any required installation, training, commissioning and other services. The evaluation of prices shall not take into account custom duties and other taxes levied on imported goods quoted CIP and sales and similar taxes levied in connection with the sale or delivery of goods. |
| Abnormally Low Bids | 39.1 An Abnormally Low Bid is one where the Bid price, in combination with other elements of the Bid, appears so low that it raises material concerns with the Purchaser as to the capability of the Bidder to perform the Contract for the offered Bid Price. ### 39.2 In the event of identification of a potentially Abnormally Low Bid, the Purchaser shall seek written clarification from the Bidder, including a detailed price analyses of its Bid price in relation to the subject matter of the contract, scope, delivery schedule, allocation of risks and responsibilities and any other requirements of the bidding document. 39.3 After evaluation of the price analyses, in the event that the Purchaser determines that the Bidder has failed to demonstrate its capability to perform the contract for the offered Bid price, the Purchaser shall reject the Bid. |
| Evaluation of Combined Technical and Financial Parts, Most Advantageous Bid and Notification of Intention to Award |  |
| Evaluation of combined Technical and Financial Parts | - The Purchaser’s evaluation of responsive Bids will take into account technical factors, in addition to cost factors in accordance with Section III Evaluation and Qualification Criteria. The weight to be assigned for the Technical factors and cost is specified in the BDS. The Purchaser will rank the Bids based on the evaluated Bid score (B). - The Purchaser will determine the Most Advantageous Bid. The Most Advantageous Bid is the Bid of the Bidder that meets the Qualification Criteria and whose Bid has been determined to be substantially responsive to the Bidding document and is the Bid with the highest combined technical and financial score. |
| Purchaser’s Right to Accept Any Bid, and to Reject Any or All Bids | - The Purchaser reserves the right to accept or reject any Bid, and to annul the Bidding process and reject all Bids at any time prior to Contract Award, without thereby incurring any liability to Bidders. In case of annulment, all Bids submitted and specifically, Bid securities, shall be promptly returned to the Bidders. |
| Standstill Period | - The Contract shall not be awarded earlier than the expiry of the Standstill Period. The Standstill Period shall be ten (10) Business Days unless extended in accordance with ITB 47. The Standstill Period commences the day after the date the Purchaser has transmitted to each Bidder the Notification of Intention to Award the Contract. Where only one Bid is submitted, or if this contract is in response to an emergency situation recognized by the Bank, the Standstill Period shall not apply. |
| Notification of Intention to Award | 43.1 The Purchaser shall send to each Bidder (that has not already been notified that it has been unsuccessful) the Notification of Intention to Award the Contract to the successful Bidder. The Notification of Intention to Award shall contain, at a minimum, the following information: - the name and address of the Bidder submitting the successful Bid; - the Contract price of the successful Bid; - the total combined score of the successful Bidder; - the names of all Bidders who submitted Bids, and their Bid prices as readout, and as evaluated and technical scores; - a statement of the reason(s) the Bid (of the unsuccessful Bidder to whom the notification is addressed) was unsuccessful; - the expiry date of the Standstill Period; - instructions on how to request a debriefing and/or submit a complaint during the standstill period. |
| Award of Contract |  |
| Award Criteria | - Subject to ITB 41, the Purchaser shall award the Contract to the successful Bidder. This is the Bidder whose Bid has been determined to be the Most Advantageous Bid as specified in ITB 40. |
| Purchaser’s Right to Vary Quantities at Time of Award | - At the time the Contract is awarded, the Purchaser reserves the right to increase or decrease the quantity of Goods and Related Services originally specified in Section VII, Schedule of Requirements, provided this does not exceed the percentages specified in the BDS, and without any change in the unit prices or other terms and conditions of the Bid and the bidding document. |
| Notification of Award | 46.1 Prior to the date of expiry of the Bid validity and upon expiry of the Standstill Period, specified in ITB 42.1 or any extension thereof, and upon satisfactorily addressing any complaint that has been filed within the Standstill Period, the Purchaser shall notify the successful Bidder, in writing, that its Bid has been accepted. The notification of award (hereinafter and in the Contract Forms called the “Letter of Acceptance”) shall specify the sum that the Purchaser will pay the Supplier in consideration of the execution of the Contract (hereinafter and in the Conditions of Contract and Contract Forms called “the Contract Price”). 46.2 Within ten (10) Business Days after the date of transmission of the Letter of Acceptance, the Purchaser shall publish the Contract Award Notice which shall contain, at a minimum, the following information: - name and address of the Purchaser; - name and reference number of the contract being awarded, and the selection method used; - names of all Bidders that submitted Bids, and their Bid prices as read out at Bid opening, and as evaluated; - names of all Bidders whose Bids were rejected either as nonresponsive or as not meeting qualification criteria, or were not evaluated, with the reasons therefor; - the name of the successful Bidder, the final total contract price, the contract duration and a summary of its scope; and - successful Bidder’s Beneficial Ownership Disclosure Form. 46.3 The Contract Award Notice shall be published on the Purchaser’s website with free access if available, or in at least one newspaper of national circulation in the Purchaser’s Country, or in the official gazette. The Purchaser shall also publish the contract award notice in UNDB online. 46.4 Until a formal Contract is prepared and executed, the Letter of Acceptance shall constitute a binding Contract. |
| Debriefing by the Purchaser | - On receipt of the Purchaser’s Notification of Intention to Award referred to in ITB 43.1, an unsuccessful Bidder has three (3) Business Days to make a written request to the Purchaser for a debriefing. The Purchaser shall provide a debriefing to all unsuccessful Bidders whose request is received within this deadline. - Where a request for debriefing is received within the deadline, the Purchaser shall provide a debriefing within five (5) Business Days, unless the Purchaser decides, for justifiable reasons, to provide the debriefing outside this timeframe. In that case, the standstill period shall automatically be extended until five (5) Business Days after such debriefing is provided. If more than one debriefing is so delayed, the standstill period shall not end earlier than five (5) Business Days after the last debriefing takes place. The Purchaser shall promptly inform, by the quickest means available, all Bidders of the extended standstill period - Where a request for debriefing is received by the Purchaser later than the three (3) Business Day deadline, the Purchaser should provide the debriefing as soon as practicable, and normally no later than fifteen (15) Business Days from the date of publication of Public Notice of Award of contract. Requests for debriefing received outside the three (3) day deadline shall not lead to extension of the standstill period. - Debriefings of unsuccessful Bidders may be done in writing or verbally. The Bidders shall bear their own costs of attending such a debriefing meeting. |
| Signing of Contract | 48.1 The Purchaser shall send to the successful Bidder the Letter of Acceptance including the Contract Agreement, and a request to submit the Beneficial Ownership Disclosure Form providing additional information on its beneficial ownership. The Beneficial Ownership Disclosure Form shall be submitted within eight (8) Business Days of receiving this request. 48.2 The successful Bidder shall sign, date and return to the Purchaser, the Contract Agreement within twenty-eight (28) days of its receipt. 48.3 Notwithstanding ITB 48.2 above, in case signing of the Contract Agreement is prevented by any export restrictions attributable to the Purchaser, to the country of the Purchaser, or to the use of the products/goods, systems or services to be supplied, where such export restrictions arise from trade regulations from a country supplying those products/goods, systems or services, the Bidder shall not be bound by its Bid, always provided however, that the Bidder can demonstrate to the satisfaction of the Purchaser and of the Bank that signing of the Contact Agreement has not been prevented by any lack of diligence on the part of the Bidder in completing any formalities, including applying for permits, authorizations and licenses necessary for the export of the products/goods, systems or services under the terms of the Contract. |
| Performance Security | - Within twenty-eight (28) days of the receipt of the Letter of Acceptance from the Purchaser, the successful Bidder, if required, shall furnish the Performance Security in accordance with the GCC 18 using for that purpose the Performance Security Form included in Section X, Contract Forms, or another Form acceptable to the Purchaser. If the Performance Security furnished by the successful Bidder is in the form of a bond, it shall be issued by a bonding or insurance company that has been determined by the successful Bidder to be acceptable to the Purchaser. A foreign institution providing a bond shall have a correspondent financial institution located in the Purchaser’s Country, unless the Purchaser has agreed in writing that a correspondent financial institution is not required. - Failure of the successful Bidder to submit the above-mentioned Performance Security or sign the Contract shall constitute sufficient grounds for the annulment of the award and forfeiture of the Bid Security. In that event the Purchaser may award the Contract to the Bidder offering the Most Advantageous Bid. |
| Procurement Related Complaint | 50.1 The procedures for making a Procurement-related Complaint are as specified in the BDS. |

| Section II - Bid Data Sheet (BDS) The following specific data for the Goods to be procured shall complement, supplement, and/or amend the provisions in the Instructions to Bidders (ITB). Whenever there is a conflict, the provisions herein shall prevail over those in ITB. |  |
| --- | --- |
| ITB Reference | A. General |
| ITB 1.1 | The reference number of the Request for Bids (RFB) is: SO-MOF-561587-GO-RFB The Purchaser is Ministry of Finance of the Federal Government of Somalia. The name of the RFB: Supply, Delivery, Installation and Commissioning of Laboratory Equipment for the National Reference Laboratory (NRL) of the Somali Bureau of Standards (SoBS) The number and identification of lots (contracts) comprising this RFB is: One (1) lot. The lot comprises the supply, delivery, installation and commissioning of all laboratory equipment, consumables, reagents and related services listed in Section VII, Schedule of Requirements, for the National Reference Laboratory (NRL) of the Somali Bureau of Standards (SoBS). Bidders shall quote for all items and quantities of the lot; partial bids shall be rejected as non-responsive. A single Contract will be awarded to the Bidder offering the Most Advantageous Bid. |
| ITB 2.1 | The Borrower is: Federal Government of Somalia (FGS) Loan or Financing Agreement amount: US$ 70 million The name of the Project is: De-Risking, Inclusion and Value Enhancement of Pastoral Economies Project (Drive). |
| ITB 4.1 | Maximum number of members in the Joint Venture (JV) shall be: 2 |
| - ITB 4.5 | A list of debarred firms and individuals is available on the Bank’s external website: http://www.worldbank.org/debarr. |
|  | B. Contents of Bidding Document |
| ITB 7.1 | For Clarification of Bid purposes only, the Purchaser’s address is: Contact Person: Procurement Specialist Address; PIU Office (DRIVE Project) Ministry of Finance, Federal Government of Somalia Shangani District, Mogadishu, Somalia, Email: bid.drive@piu.mof.gov.so, Requests for clarification should be received by the Purchaser no later than: 14 days before the deadline for bid submission |
|  | C. Preparation of Bids |
| ITB 10.1 | The language of the Bid is: English. All correspondence exchange shall be in English language. Language for translation of supporting documents and printed literature is English. |
| ITB 11.2 (i) & 11.3 (d) | The Bidder shall submit the following additional documents in its Bid: - The completed item-by-item technical compliance matrix (the “Bidder’s Response” column of the Technical Specifications in Section VII) for all Goods line items, stating “Comply” or describing any deviation, with a precise cross-reference to the manufacturer’s published technical literature (datasheet or catalogue page) for each item; - A site-readiness and interface plan identifying, for each instrument or system, the electrical, earthing, UPS/power-conditioning, HVAC/exhaust and fume-hood, gases, water/drainage, bench and floor loading, access and offloading, networking/ICT and cold-chain interface requirements; - (The installation, commissioning, FAT/SAT (where applicable) and IQ/OQ/PQ qualification and acceptance-testing methodology and plan; - A warranty, preventive-maintenance, service and spare-parts plan, including the commitment to spare-parts availability in accordance with ITB 16.4 and the service response times specified in the SCC; - A LIMS/instrument integration and data-integrity approach (interfaces, user roles, audit trails, data ownership, backup and integration responsibilities); - For chemicals, reagents, reference standards and gases: Safety Data Sheets (SDS), certificates of analysis (where applicable), and transport, storage and cold-chain arrangements; - A detailed delivery, installation and commissioning schedule and logistics plan, including customs/import support, security and journey-management arrangements, and risk-mitigation arrangements for delivery to Mogadishu, meeting the timelines stated in Section VII; and - The compliance and standards documents listed in the Schedule of Requirements under Training and Documentation; and - A signed ESHS/OHS commitment statement in accordance with Section III (Technical Evaluation, mandatory requirement V), confirming that the Bidder will prepare and implement the contract-specific ESHS/OHS Management Plan before mobilization, comply with the DRIVE Project E&S instruments, enforce the worker Code of Conduct (including SEA/SH prevention) and notify and report incidents in accordance with the SCC. |
| ITB 13.1 | Alternative Bids (Technical and Financial Parts) shall not be considered. |
| ITB 14.5 | The prices quoted by the Bidder shall not be subject to adjustment during the performance of the Contract. |
| ITB 14.6 | Prices quoted for the bid (contract) shall correspond to 100% percent of the items specified for the lot (contract). Prices quoted for each item of a lot shall correspond to 100% percent of the quantities specified for this item of a lot. |
| ITB 14.7 | The Incoterms edition is: Incoterms 2020 |
| ITB 14.8 (a)(i) | The Incoterm for Goods manufactured outside the Purchaser’s Country shall be: DDP (Incoterms 2020), Named place of destination: Somali Bureau of Standards (SoBS) – National Reference Laboratory (NRL), Mogadishu, Federal Republic of Somalia. The quoted price shall include all costs up to delivery DDP at the Final Destination, including customs clearance, import duties and taxes; the Purchaser shall provide reasonable assistance with any available duty-exemption formalities. |
| ITB 14.8 (a)(iii), (b)(ii) and (c)(v) | Final Destination (Project Site): Somali Bureau of Standards (SoBS) – National Reference Laboratory (NRL), Mogadishu, Federal Republic of Somalia. The exact delivery address and site details shall be confirmed at contract signature. |
| ITB 14.8 (b)(i) | Place of Destination: Somali Bureau of Standards (SoBS) – National Reference Laboratory (NRL), Mogadishu, Federal Republic of Somalia. |
| ITB 15.1 | A Bidder shall quote its price entirely and shall be paid in USD only |
| ITB 16.4 | Period the Goods are expected to be functioning (for the purpose of spare parts): five (5) years from Operational Acceptance of the relevant Goods (as defined in the SCC) |
| ITB 17.2 (a) | Manufacturer’s authorization is: required – for the principal instruments, as a minimum the following line items of Section VII, Schedule of Requirements: 1 (LC-MS/MS System), 2 (GC-MS/MS System), 3 (HPLC System), 4 (GC-MS System), 5 (AAS – Flame + Graphite Furnace), 6 (FTIR Spectrometer), 9 (Ion Chromatograph), 24 (Biosafety Cabinet), 25 (Autoclave / Steam Sterilizer), 26 (Anaerobic Incubator / Workstation), 28 (Real-Time PCR System), 31 (Automated Microplate Reader – ELISA), 37 (Analytical Balance), 38 (Precision Balance), 44 (High-Speed Refrigerated Centrifuge), 63 and 64 (UPS Systems), 66 (HVAC – Dedicated Lab Precision Air Handling Unit), 85 (Environmental Monitoring System), 90 (Laboratory Server / NAS Backup System), 98 (Milk Analyzer), 99 (Hematology Analyzer) and 100 (Veterinary Biochemistry Analyzer). |
| ITB 17.2 (b) | After sales service is: required |
| ITB 18.1 | The Bid shall be valid until: 150 days from the Bid Submission deadline date |
| ITB 18.3 (a) | The Bid price shall be adjusted by the following factor(s): None |
| ITB 19.1 | A Bid-Securing Declaration shall be required. |
| ITB 19.3 (d) | Other types of acceptable securities: None |
| ITB 19.9 | If the Bidder performs any of the actions prescribed in subparagraphs ITB 19.9 (a) or (b), the Borrower will declare the Bidder ineligible to be awarded contracts by the Purchaser for a period of two (2) years starting from the date the Bidder performs any of the actions. |
| ITB 20.3 | The written confirmation of authorization to sign on behalf of the Bidder shall consist of: Notarized power of attorney |
|  | D. Submission of Bids |
| ITB 21.2 | In addition to the original of the Bid, the number of copies is: Technical Part of the Bid – One Original and 2 copies Financial Part of the Bid: One Original An electronic copy of the technical Bid only (on a generally used portable device such as flash memory device) must be included in the envelope containing the technical original hard copy. This is required for information purposes only and will not be used for evaluation. Bidders are advised not to include any financial pricing details for the Bid in either the hard or soft copy of the technical bid. Submission of such information will result in the proposal being deemed non-responsive. |
| ITB 22.1 | For Bid submission purposes only, the Purchaser’s address is: Attention: Procurement Specialist Street Address: Ministry of Finance, Opposite Central Bank of Somalia: Floor/ Room number: 6th floor, Dahab Apartments, Next to Ministry of Finance Building PIU Office (DRIVE Project) City: Mogadishu. ZIP/Postal Code: N/A Country: Federal Republic of Somalia. The deadline for Bid submission is: Date: 13th December 2026 Time: 11:00 A.M. Somalia local time Bidders shall not have the option of submitting their Bids electronically. |
|  | E. Public Opening of Technical Parts of Bids |
| ITB 25.1 | The Bid opening shall take place at: Street Address: Ministry of finance building, Opposite Central Bank of Somalia (CBS), Shangani District, Mogadishu, Somalia. Floor/Room number: Conference room, Dahab apartments, 6th floor. City: Mogadishu Country: Somalia Date: 13th December 2026 Time: 11:00 AM. Somalia time. |
| ITB 25.6 | The Letter of Bid - Technical Part and the sealed envelope marked “Second Envelope - Financial Part” shall be initialed by at least three (3) representatives of the Purchaser conducting the Bid opening. Each Bid shall be numbered, and any modification to the Bid shall be initialed by the Purchaser’s representatives. |
| G. Evaluation of Technical Parts of Bids |  |
| ITB 32.4 | The technical factors and sub factors as applicable and the corresponding scores out of 100% are: No.; Category; Weight. I; Compliance and performance of the offered equipment with the Technical Specifications; 40%. II; Installation, commissioning and qualification (IQ/OQ/PQ) methodology and implementation plan; 20%. III; After-sales support, warranty, spare parts and maintenance; 15%. IV; Training, knowledge transfer and documentation; 15%. V; Delivery schedule and logistics; 10%. Total; 100%. Mandatory technical items will be excluded from technical scoring. Bids which do not satisfy all Mandatory requirements will be excluded from further consideration. A Bid that does not meet all mandatory requirements listed in Section III and Section VII shall be rejected as non-responsive before scored technical evaluation; mandatory minimum compliance shall not be scored. Technical scoring will apply to desirable features and sub-features as listed in the tables for each category below. The evidence required, the full-score benchmarks and the lower-score / zero-score guidance for each sub-feature are specified in Section III, Evaluation and Qualification Criteria. Each sub-feature will be given a score between 0 and 6 as follows: - 0 for the sub-feature being absent; - 2 for the sub-feature being present but showing deficiencies; - 4 for meeting the requirements; - 6 for marginally exceeding the requirements; As described for each Category below, the Category technical score will be calculated by combining the sub-feature scores as weighted sum. The five categories of technical scores will be combined based on the weights for each Category in the table above. Category I: Compliance and performance of the offered equipment with the Technical Specifications – 40 No; Sub-feature; Score. 1.1; Completeness and technical conformity of the offered equipment with the specifications in Section VII, supported by manufacturer datasheets, brochures and standards certificates; 20. 1.2; Performance characteristics and quality of the principal instruments (sensitivity, resolution, accuracy, throughput, automation and upgradeability), including software compliance (GLP/GMP, audit trail, 21 CFR Part 11-ready where applicable); 12. 1.3; Certification of equipment to international standards (e.g. CE marking, EN/ISO/NSF/IVD as applicable per item) and suitability for use in an ISO/IEC 17025 accredited laboratory; 8. Total score for category 1; 40 points. Category II: Installation, commissioning and qualification (IQ/OQ/PQ) methodology and implementation plan - 20 No; Sub-feature; Score. 2.1; Detailed installation and commissioning methodology, including site readiness assessment, pre-installation checklist, utility/environmental requirements and unpacking/positioning plan; 8. 2.2; Installation/Operational/Performance Qualification (IQ/OQ/PQ) protocols and acceptance testing procedures with traceable documentation and calibration certificates on completion; 8. 2.3; Qualifications and experience of the proposed installation and commissioning engineers (CVs and manufacturer factory certification); 4. Total score for category 2; 20 points. Category III: After-sales support, warranty, spare parts and maintenance – 15 No; Sub-feature; Score. 3.1; Warranty coverage (minimum twenty-four (24) months from Operational Acceptance, or thirty (30) months from delivery, whichever occurs first), its scope, service response times, loan/replacement and escalation arrangements for critical downtime, and commitment to minimum uptime and timely support during ownership and operation of the equipment; 6. 3.2; Written assurance on availability of spare parts and consumables for at least five (5) years from Operational Acceptance (ITB 16.4); 5. 3.3; Preventive maintenance plan and service arrangements covering the warranty period (local or regional service presence, remote diagnostics, spare-parts kits); 4. Total score for category 3; 15 points. Category IV: Training, knowledge transfer and documentation - 15 No; Sub-feature; Score. 4.1; Detailed description of how the bidder will undertake effective training, knowledge transfer, handholding and capacity building for SoBS analysts over the implementation and warranty periods (curriculum, duration, number of trainees); 8. 4.2; Supply of all brochures, manuals, SOPs and training material in English, along with relevant documentary evidence stated in the bidding document; 7. Total score for category 4; 15 points. Category V: Delivery schedule and logistics - 10 No; Sub-feature; Score. 5.1; Detailed delivery, installation and commissioning schedule that meets the timelines stated in Section VII, including logistics, customs clearance and risk-mitigation arrangements for delivery to Mogadishu; 10. Total score for category 5; 10 points. |
| H. Notification of Evaluation of Technical Parts and Public Opening of Financial Parts of Bids |  |
| ITB 33.5 | The Letter of Bid – Financial Part and the Price Schedules shall be initialed by at least three (3) representatives of the Purchaser conducting the Financial Part opening. Each Financial Part shall be numbered, and any modification to the unit or total price shall be initialed by the Purchaser’s representatives. |
|  | I. Evaluation of Financial Part of Bids |
| ITB 34.2(a) | Bids will be evaluated for all items of the single lot together. Bidders shall quote for all items and quantities listed in Section VII, Schedule of Requirements; incomplete bids shall be rejected as nonresponsive. The Contract will comprise all the Goods and Related Services awarded to the successful Bidder offering the Most Advantageous Bid. |
| ITB 34.6 | The adjustments shall be determined using the following criteria, from amongst those set out in Section III, Evaluation and Qualification Criteria: - Deviation in Delivery schedule: No monetary adjustment; delivery beyond the latest delivery/completion dates specified in Section VII, Schedule of Requirements shall be treated as nonresponsive. - Deviation payment schedule: No. - the cost of major replacement component, mandatory spare parts, and service: No separate monetary adjustment; Bidders shall provide the required five-year spare-parts and service commitments in accordance with ITB 16.4 and Section VII. - the availability in the Purchaser’s Country of spare parts and after-sales services for the equipment offered in the Bid No adjustments but The supplier must confirm ability to provide required after sales support such as replacement of defective item including supply of spare parts where required - Life cycle costs: the costs during the life of the goods or equipment No. - the performance and productivity of the equipment offered; No. |
| ITB 36.1 | Not Applicable. |
| ITB 37.1 | A margin of domestic preference shall not apply. |
| J. Evaluation of Combined Technical and Financial Parts and Most Advantageous Bid |  |
| ITB 40.1 | The weight to be given for cost is: 0.3 |
|  | J. Award of Contract |
| ITB 45.1 | The maximum percentage by which quantities may be increased is: 15% The maximum percentage by which quantities may be decreased is: 15%. Any variation shall not alter the unit prices or other terms and conditions of the Bid and the Contract. |
| ITB 50.1 | The procedures for making a Procurement-related Complaint are detailed in the “Procurement Regulations for IPF Borrowers (Annex III).” If a Bidder wishes to make a Procurement-related Complaint, the Bidder should submit its complaint following these procedures, in writing (by the quickest means available, that is either by email or fax), to: For the attention: Saleiman S. Omar Title/position: Director General Purchaser: Ministry of Finance Email address: dg@mof.gov.so, A copy of the complaint can be sent for the Bank’s information and monitoring to: pprocurementcomplaints@worldbank.org . In summary, a Procurement-related Complaint may challenge any of the following: - the terms of the Bidding Documents; - the Purchaser’s decision to exclude a Bidder from the procurement process prior to the award of contract; and - the Purchaser’s decision to award the contract. |

Section III - Evaluation and Qualification Criteria

This Section contains the criteria that the Purchaser shall use to evaluate Bids and qualify the Bidders.  No other factors, methods or criteria shall be used other than specified in this bidding document. The Proposer shall provide all the information requested in the forms included in Section IV, Proposal Forms.

TECHNICAL PART

Qualification

- 1.1	Eligibility criteria

- The Proposer’s eligibility shall be assessed in accordance with the preliminary evaluation table included in this section.

| Preliminary Evaluation |  | Met / Not Met |
| --- | --- | --- |
| 1 | If the proposer is not the owner of the solution to be implemented a valid Manufacturer Authorization Form (MAF), issued by OEM should be provided. |  |
| 2 | Valid registration certificate issued by any legally recognised jurisdiction for the proposer and any associated joint venture parties. |  |
| 3 | A signed Power of Attorney. If the person signing the proposal is the Chief Executive Officer (CEO), Managing Director (MD) or Owner of the bidding firm, a power of attorney shall not be required. |  |
| 4 | Duly signed and stamped proposal-securing declaration. |  |
| 5 | Latest 3 year audited financial reports 2023, 2024 and 2025. |  |
| 6 | Joint venture agreement contract or letter of intent which should include capacity/role and responsibility of each firm in the project. |  |
| 7 | Duly filled and signed Letter of Proposal – Technical Part |  |

1.2 Qualification Criteria (ITB 32.1)

The Purchaser shall assess each Bid against the following Qualification Criteria. Requirements not included in the text below shall not be used in the evaluation of the Bidder’s qualifications.

- Financial Capability: The Bidder shall submit audited financial statements or, if not required by the law of the Bidder’s country, other financial statements acceptable to the Purchaser, for the last Three (3) years prior to bid submission deadline, demonstrating the current soundness of the Bidder’s financial position. For a joint venture, this requirement shall be met by each member. In addition, the Bidder shall demonstrate a minimum average annual turnover in the last three (3) years (2023, 2024 and 2025) of not less than USD 3,000,000 (Three million United States Dollars). For a joint venture, the turnover requirement shall be met by all members combined. In addition, the Bidder shall demonstrate access to liquid assets, lines of credit or other financial resources, net of other contractual commitments, of not less than USD 750,000 (seven hundred and fifty thousand United States Dollars). For a joint venture, the monetary requirements of this sub-criterion may be met by all members combined, provided that the lead member meets at least forty percent (40%) of each monetary requirement and each other member meets at least twenty-five percent (25%) of each monetary requirement;

- Specific Experience: The Bidder shall demonstrate that it has successfully completed at least two (2) contracts within the last five (5) years prior to bid submission deadline, involving goods of similar nature and complexity (supply, delivery, installation and commissioning of laboratory, analytical, scientific or medical equipment), each with a minimum value of USD 2,500,000 (Two million Five Hundred Thousand United States Dollars). The contracts shall have been successfully and substantially completed and shall be similar in nature and complexity to the Goods and Related Services under the Contract. Similar contracts shall involve the integrated supply of laboratory, analytical, scientific or medical equipment, including installation, commissioning, calibration/verification, training and multi-OEM after-sales support.

- Documentary Evidence: The Bidder shall furnish documentary evidence to demonstrate that the Goods it offers meet the following usage requirement:

- Detailed technical literature (catalogues, brochures and manufacturer’s data sheets) for each offered item, cross-referenced item-by-item to the Technical Specifications in Section VII, Schedule of Requirements, demonstrating full conformity with the stated requirements.

- Evidence that the offered equipment is suitable for use in a laboratory operating, or intending to operate, in accordance with ISO/IEC 17025, including, where applicable, instrument software features supporting GLP/GMP and data-integrity requirements (audit trails, user access control).

- Evidence that the offered equipment carries relevant international product certification (e.g. CE marking, and conformity with applicable EN/ISO/IEC standards; IVD/IVDR compliance for in-vitro diagnostic instruments where applicable).

- Detailed after-sales support and services plan covering installation, commissioning, IQ/OQ/PQ qualification, warranty-period support, response times and a commitment to achieve the required minimum instrument up-time during ownership and operation of the proposed equipment.

- Manufacturer’s quality management certification (ISO 9001 or equivalent) covering the design and manufacture of the principal instruments offered.

- Written assurance on availability of spare parts and consumables for at least five (5) years from Operational Acceptance.

- Detailed description of how the bidder will undertake effective training, knowledge transfer, handholding and capacity building over the implementation and warranty periods.

- Supply of all brochures, manuals, SOPs and training material in English.

- Detailed delivery, installation and commissioning schedule that meets the timelines stated in Section VII, Schedule of Requirements.

- Completed item-by-item technical compliance matrix (the “Bidder’s Response” column of the Technical Specifications in Section VII) for all Goods line items, in accordance with BDS ITB 11.2 (i) & 11.3 (d).

- Site-readiness and interface plan identifying, for each instrument or system, the electrical, earthing, UPS/power-conditioning, HVAC/exhaust and fume-hood, gases, water/drainage, bench and floor loading, access and offloading, networking/ICT and cold-chain interface requirements.

- For chemicals, reagents, reference standards and gases: Safety Data Sheets (SDS), certificates of analysis (where applicable), and transport, storage and cold-chain arrangements.

- LIMS/instrument integration and data-integrity approach, including interfaces, user roles, audit trails, data ownership, backup and integration responsibilities.

- Manufacturing experience and Technical Capacity: For the items under the Contract that the bidder is a manufacturer, the Bidder shall furnish documentary evidence to demonstrate that:

- it has manufactured goods of similar nature and complexity for at least Five (5) years, prior to the bid submission deadline; and

- it maintains a documented quality management system certified to ISO 9001 (or equivalent), covering the design and manufacture of the goods offered.

- Manufacturer’s authorization: A Bidder who does not manufacture an item/s where a manufacturer authorization is required in accordance with BDS ITB 17.2 (a), the Bidder shall provide evidence of being duly authorized by a manufacturer (Manufacturer’s Authorization Form, Section IV, Bidding Forms), meeting the criteria in (d) (i) and (ii) above, to supply the Goods.

- A bidder who does not manufacture an item/s where a manufacturer authorization is not required in accordance with BDS ITB 17.2 (a), the bidder shall submit documentation on, its status as a supplier, to the satisfaction of the Purchaser (e.g. authorized dealer/ distributor of the items).

At the time of Contract Award, the Bidder (including each subcontractor proposed by the Bidder) shall not be subject to disqualification by the Bank for non-compliance with SEA/ SH obligations.

Technical Evaluation (ITB 32.4)

Assessment of adequacy of Technical Part with the requirements:

| Ref | Requirement Description | Met/ Not met |
| --- | --- | --- |
| I | The Bidder must have successfully supplied, installed and commissioned similar laboratory, analytical, scientific or medical equipment to at least two (2) institutions (e.g. national reference laboratories, food and feed testing laboratories, or similar), supported by client reference letters or completion certificates. |  |
| II | The Bidder must be the Original Equipment Manufacturer (OEM) of the principal instruments offered, or an authorized agent/distributor demonstrating direct linkage to the OEM, supported by the Manufacturer’s Authorization Form (Section IV, Bidding Forms). |  |
| III | Each offered item must be accompanied by complete technical literature (catalogues, brochures and data sheets) cross-referenced to the Technical Specifications in Section VII, demonstrating full conformity; the offered goods must be new, unused, of the most recent or current models, and incorporate all recent improvements in design and materials. |  |
| IV | The Bidder must demonstrate after-sales support capability: written assurance of availability of spare parts and consumables for at least five (5) years, and a service and maintenance plan covering installation, commissioning, IQ/OQ/PQ qualification and the warranty period. |  |
| V | The Bidder must submit a signed ESHS/OHS commitment statement confirming that it will: (i) prepare and implement a contract-specific ESHS/OHS Management Plan acceptable to the Purchaser before mobilization to site; (ii) comply with the World Bank Environmental and Social Standards applicable to the Project and the DRIVE Project E&S instruments (ESCP, ESMF, SEP and the E&S procedures of the Project Implementation Manual); (iii) enforce the worker Code of Conduct, including SEA/SH prevention, for all personnel engaged on site; and (iv) notify and report incidents in accordance with the SCC. |  |

Mandatory pass/fail requirements listed in this Section III and in Section VII are not scored. A Bid that does not meet all mandatory requirements shall be rejected as nonresponsive before scored technical evaluation. Only Bids that meet all mandatory requirements shall be scored using the Rated Criteria specified in BDS ITB 32.4 and the scoring methodology below.

Technical Part Scoring Methodology

Each sub-feature will be given a score between 0 and 6 as follows:

| Score (of the total score for the factor/subfactor as applicable | Description | Remarks |
| --- | --- | --- |
| 0 | The required sub-feature being absent; |  |
| 2 | Required sub-feature being present but showing deficiencies; |  |
| 4 | Sufficient information to demonstrate how stated requirement will be met |  |
| 6 | Sufficient information to demonstrate that the requirement will be marginally exceeded |  |

| No. | Category | Weight |
| --- | --- | --- |
| I | Compliance and performance of the offered equipment with the Technical Specifications | 40% |
| II | Installation, commissioning and qualification (IQ/OQ/PQ) methodology and implementation plan | 20% |
| III | After-sales support, warranty, spare parts and maintenance | 15% |
| IV | Training, knowledge transfer and documentation | 15% |
| V | Delivery schedule and logistics | 10% |
| Total |  | 100% |

Category I: Compliance and performance of the offered equipment with the Technical Specifications - 40

| No | Sub-feature | Score |
| --- | --- | --- |
| 1.1 | Completeness and technical conformity of the offered equipment with the specifications in Section VII, supported by manufacturer datasheets, brochures and standards certificates | 20 |
| 1.2 | Performance characteristics and quality of the principal instruments (sensitivity, resolution, accuracy, throughput, automation and upgradeability), including software compliance (GLP/GMP, audit trail, 21 CFR Part 11-ready where applicable) | 12 |
| 1.3 | Certification of equipment to international standards (e.g. CE marking, EN/ISO/NSF/IVD as applicable per item) and suitability for use in an ISO/IEC 17025 accredited laboratory | 8 |
| Total score for category 1 |  | 40 points |

Category II: Installation, commissioning and qualification (IQ/OQ/PQ) methodology and implementation plan - 20

| No | Sub-feature | Score |
| --- | --- | --- |
| 2.1 | Detailed installation and commissioning methodology, including site readiness assessment, pre-installation checklist, utility/environmental requirements and unpacking/positioning plan | 8 |
| 2.2 | Installation/Operational/Performance Qualification (IQ/OQ/PQ) protocols and acceptance testing procedures with traceable documentation and calibration certificates on completion | 8 |
| 2.3 | Qualifications and experience of the proposed installation and commissioning engineers (CVs and manufacturer factory certification) | 4 |
| Total score for category 2 |  | 20 points |

Category III: After-sales support, warranty, spare parts and maintenance - 15

| No | Sub-feature | Score |
| --- | --- | --- |
| 3.1 | Warranty coverage (minimum twenty-four (24) months from Operational Acceptance, or thirty (30) months from delivery, whichever occurs first), its scope, service response times, loan/replacement and escalation arrangements for critical downtime, and commitment to minimum uptime and timely support during ownership and operation of the equipment | 6 |
| 3.2 | Written assurance on availability of spare parts and consumables for at least five (5) years from Operational Acceptance (ITB 16.4) | 5 |
| 3.3 | Preventive maintenance plan and service arrangements covering the warranty period (local or regional service presence, remote diagnostics, spare-parts kits) | 4 |
| Total score for category 3 |  | 15 points |

Category IV: Training, knowledge transfer and documentation - 15

| No | Sub-feature | Score |
| --- | --- | --- |
| 4.1 | Detailed description of how the bidder will undertake effective training, knowledge transfer, handholding and capacity building for SoBS analysts over the implementation and warranty periods (curriculum, duration, number of trainees) | 8 |
| 4.2 | Supply of all brochures, manuals, SOPs and training material in English, along with relevant documentary evidence stated in the bidding document | 7 |
| Total score for category 4 |  | 15 points |

Category V: Delivery schedule and logistics - 10

| No | Sub-feature | Score |
| --- | --- | --- |
| 5.1 | Detailed delivery, installation and commissioning schedule that meets the timelines stated in Section VII, including logistics, customs clearance and risk-mitigation arrangements for delivery to Mogadishu | 10 |
| Total score for category 5 |  | 10 points |

The score for each sub- factor (i) within a factor (j) will be combined with the scores of sub- factors in the same factor as a weighted sum to form the Factor Technical Score using the following formula:

where:

tji	= the technical score for sub- factor “i” in factor “j”,

wji	= the weight of sub- factor “i” in factor “j”,

k	= the number of scored sub-factors in factor “j”, and

The Factor Technical Scores will be combined in a weighted sum to form the total Technical Bid Score using the following formula:

where:

Sj	= the Factor Technical Score of factor “j”,

Wj	= the weight of factor “j” as specified in the BDS,

n	= the number of Factors, and

Rated Criteria – evidence, benchmarks and scoring guidance. The following table specifies, for each factor and sub-feature of BDS ITB 32.4, the evidence to be submitted in the Technical Part, the benchmark for award of the full sub-feature score, and guidance for lower or zero scores. Scores of 0, 2, 4 or 6 shall be awarded for each sub-feature in accordance with the methodology above, applied consistently with this guidance.

| Factor / sub-feature (maximum points) | Evidence required in the Technical Part | Full-score benchmark | Lower-score / zero-score guidance |
| --- | --- | --- | --- |
| Category I – Compliance and performance of offered equipment with the Technical Specifications (40 points) |  |  |  |
| 1.1 Completeness and technical conformity (20) | Completed item-by-item compliance matrix for all Goods line items of Section VII, with manufacturer datasheets, model numbers, configuration lists and standards certificates. | All items are addressed with specific make/model/configuration and unambiguous evidence of compliance or superior performance. | Lower scores for generic catalogues, unclear configuration, missing items or unsupported equivalence. Zero if no usable evidence. |
| 1.2 Performance and quality of principal instruments, incl. software compliance (12) | Published performance data for the principal instruments (sensitivity, resolution, accuracy, throughput, automation, upgradeability) and software data-integrity features (GLP/GMP, audit trail, 21 CFR Part 11-ready where applicable). | Documented performance meeting or exceeding Section VII for all principal instruments, with demonstrated data-integrity features. | Lower scores for partial coverage or marginal/unsupported performance claims. |
| 1.3 International certification and ISO/IEC 17025 suitability (8) | Valid CE/EN/ISO/NSF/IVD certificates, as applicable per item, and evidence of suitability for use in an ISO/IEC 17025 laboratory. | Valid certification for all principal instruments and clear 17025-suitability evidence. | Lower scores where certificates are missing for some items. Zero if no certification evidence. |
| Category II – Installation, commissioning and qualification (IQ/OQ/PQ) methodology and implementation plan (20 points) |  |  |  |
| 2.1 Installation methodology and site readiness (8) | Site-readiness/interface plan, pre-installation checklist, utility/environmental requirements and unpacking/positioning plan. | Item-level interface matrix (electrical, earthing, UPS, HVAC/exhaust, gases, water/drainage, loads, access, networking, cold-chain) with a realistic verification method. | Lower scores for a generic, non-itemised plan. Zero if absent. |
| 2.2 IQ/OQ/PQ protocols and acceptance testing (8) | FAT/SAT (where applicable) and IQ/OQ/PQ protocols, acceptance criteria, and calibration/verification through ISO/IEC 17025-accredited or equivalent laboratories. | Complete, traceable protocols with acceptance criteria per instrument group and SI-traceable calibration/verification. | Lower scores for incomplete or non-instrument-specific protocols. |
| 2.3 Qualifications of installation/commissioning engineers (4) | CVs and manufacturer factory certification of the proposed engineers. | Factory-certified engineers for all principal instruments. | Lower scores for uncertified or insufficient staffing. |
| Category III – After-sales support, warranty, spare parts and maintenance (15 points) |  |  |  |
| 3.1 Warranty coverage, scope and response times (6) | Warranty terms meeting the minimum of twenty-four (24) months from Operational Acceptance (or thirty (30) months from delivery, whichever occurs first), scope, response-time commitments, uptime commitment, and loan/escalation arrangements for critical downtime. | Meets or exceeds the minimum with binding response times and a loan/replacement commitment for critical downtime. | Lower scores where the minimum is met but service levels are weak. Zero if below the minimum warranty. |
| 3.2 Spare-parts and consumables availability (5) | Written, OEM-backed assurance of availability of spare parts and consumables for at least five (5) years from Operational Acceptance. | Assurance covering all principal instruments. | Lower scores for partial coverage. Zero if absent. |
| 3.3 Preventive maintenance and service arrangements (4) | Preventive-maintenance plan and service arrangements (local or regional service presence, remote diagnostics, spare-parts kits). | Credible local or regional service arrangement with defined response logistics for Somalia. | Lower scores for remote-only support with long mobilization times. |
| Category IV – Training, knowledge transfer and documentation (15 points) |  |  |  |
| 4.1 Training plan and knowledge transfer (8) | Training plan covering the Related Services courses: curriculum, duration, number of trainees, trainer qualifications, competency checks and post-training support. | Meets Section VII training requirements with clear curricula, qualified trainers and post-training support. | Lower scores for a generic outline. Zero if absent. |
| 4.2 Manuals, SOPs and documentation package (7) | Manuals, SOPs and training materials in English, and the handover documentation package (certificates, warranty documents, software licenses and administrator credentials, as-installed records). | Complete documentation package including as-installed records and administrator credentials. | Lower scores for an incomplete package. |
| Category V – Delivery schedule and logistics (10 points) |  |  |  |
| 5.1 Delivery, installation and commissioning schedule and logistics (10) | Integrated delivery/installation/commissioning schedule meeting the timelines in Section VII, with logistics, customs/import support and risk-mitigation arrangements for delivery to Mogadishu. | Realistic, itemised schedule within the latest delivery dates, with credible logistics and risk mitigation for Somalia. | Lower scores for unrealistic or incomplete schedules. Bids offering delivery beyond the latest dates are nonresponsive. |

FINANCIAL PART

Margin of Preference (ITB 37) N/A

If the Bidding Data Sheet so specifies, the Purchaser will grant a margin of preference to goods manufactured in the Purchaser’s country for the purpose of Bid comparison, in accordance with the procedures outlined in subsequent paragraphs.

Substantially responsive Bids will be classified in one of three groups, as follows:

(a)	Group A: Bids offering goods manufactured in the Purchaser’s Country, for which (i) labor, raw materials, and components from within the Purchaser’s Country account for more than thirty (30) percent of the EXW price; and (ii) the production facility in which they will be manufactured or assembled has been engaged in manufacturing or assembling such goods at least since the date of Bid submission.

(b)	Group B: All other Bids offering Goods manufactured in the Purchaser’s Country.

(c)	Group C: Bids offering Goods manufactured outside the Purchaser’s Country that have been already imported or that will be imported.

To facilitate this classification by the Purchaser, the Bidder shall complete whichever version of the Price Schedule furnished in the bidding document is appropriate provided, however, that the completion of an incorrect version of the Price Schedule by the Bidder shall not result in rejection of its Bid, but merely in the Purchaser’s reclassification of the Bid into its appropriate Bid group.

The Purchaser will first review the Bids to confirm the appropriateness of, and to modify as necessary, the Bid group classification to which Bidders assigned their Bids in preparing their Bid Forms and Price Schedules.

Following the combined evaluation procedure described below, the Bids in each group will then be compared to determine the Most Advantageous Bid in that group. The Most Advantageous Bid from each group shall then be compared with each other and if as a result of this comparison a Bid from Group A or Group B is the Most Advantageous, it shall be selected for the award.

If as a result of the preceding comparison, a Bid from Group C is the Most Advantageous Bid, all Bids from Group C shall be further compared with the Bid with the lowest evaluated cost from Group A after adding to the evaluated price of goods offered in each Bid from Group C, for the purpose of this further comparison only, an amount equal to 15% (fifteen percent) of the respective CIP Bid price for goods to be imported and already imported goods. Both prices shall include unconditional discounts and be corrected for arithmetical errors. If the Bid from Group A is the Most Advantageous, it shall be selected for award. If not, the Most Advantageous Bid from Group C shall be selected.

Evaluation Criteria (ITB 34.6)

The Purchaser shall use the criteria and methodologies listed in this Section to evaluate the Financial Part.

The Purchaser’s evaluation of the Financial Part may take into account, in addition to the Bid Price, one or more of the following factors as specified in BDS ITB 34.6, using the following criteria and methodologies.

(a)	Delivery schedule. (As per Incoterms specified in the BDS) N/A

(b)	Deviation in payment schedule. N/A

(c)	Cost of major replacement components, mandatory spare parts, and service. N/A

(d)	Availability in the Purchaser’s Country of spare parts and after sales services for equipment offered in the Bid. N/A

(e)    Life Cycle Cost: N/A

(f)	Performance and productivity of the equipment. N/A

(g)	Specific additional criteria:

Combined Evaluation

The Purchaser will evaluate and compare the Bids that have been determined to be substantially responsive.

The Purchaser’s evaluation of responsive Bids will take into account technical factors, in addition to cost factors.

An Evaluated Bid Score (B) will be calculated for each responsive Bid using the following formula (for comparison in percentages), which permits a comprehensive assessment of the Bid price and the technical merits of each Bid:

where C = Evaluated Bid Price C low = the lowest of all Evaluated Bid Prices among responsive Bids T = the total Technical Score awarded to the Bid Thigh = the Technical Score achieved by the Bid that was scored best among all responsive Bids X = weight for the Cost as specified in the BDS The Bid with the best evaluated Bid Score (B) among responsive Bids shall be the Most Advantageous Bid provided the Bidder is qualified to perform the Contract.

Multiple Contracts (ITB 34.4) N/A

Alternative Bids (ITB 13.1)

An alternative if permitted under ITB 13.1, will be evaluated as follows: N/A

- Qualification

| Factor | # 1.1 Eligibility |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| ##### Sub-Factor | ##### Criteria |  |  |  |  | ##### Documentation Required |
|  | ##### Requirement | ##### Bidder |  |  |  |  |
|  |  | Single Entity | ##### Joint Venture (existing or intended) |  |  |  |
|  |  |  | All members combined | ##### Each member | At least one member |  |
| ## 1.1 Nationality | Nationality in accordance with ITP 4.4. | Must meet requirement | Must meet requirement | Must meet requirement | N / A | Form ELI –1.1 and 1.2, with attachments |
| ## 1.2 Conflict of Interest | No- conflicts of interests as described in ITP 4.2. | Must meet requirement | Must meet requirement | Must meet requirement | N / A | Letter of Proposal |
| ## 1.3 Bank Ineligibility | Not having been declared ineligible by the Bank as described in ITP 4.5. | Must meet requirement | Must meet requirement | Must meet requirement | N / A | Letter of Proposal |
| ## 1.4 State owned Entity of the Borrower country | Compliance with conditions of ITP 4.6 | Must meet requirement | Must meet requirement | Must meet requirement | N / A | Form ELI –1.1 and 1.2, with attachments |
| ## 1.5 United Nations resolution or Borrower’s country law | Not having been excluded as a result of prohibition in the Borrower’s country laws or official regulations against commercial relations with the Proposer’s country, or by an act of compliance with UN Security Council resolution, both in accordance with ITP 4.8 | Must meet requirement | Must meet requirement | Must meet requirement | N / A | Letter of Proposal |

| Factor | # 1.2 Historical Contract Non-Performance |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| ##### Sub-Factor | ##### Criteria |  |  |  |  | Documentation Required |
|  | ##### Requirement | ##### Proposer |  |  |  |  |
|  |  | Single Entity | ##### Joint Venture (existing or intended) |  |  |  |
|  |  |  | All members combined | Each member | At least one member |  |

| ## 1.2.1 History of non-performing contracts | Non-performance of a contract did not occur as a result of Proposer’s default since 1st January 2019 | Must meet requirement by itself or as member to past or existing JV | N / A | Must meet requirement | N / A | Form CON - 2 |
| --- | --- | --- | --- | --- | --- | --- |
| ## 1.2.2 Suspension | Not under suspension based on execution of a Proposal Securing Declaration or Proposal Securing Declaration pursuant to ITP 4.7 and ITP 20.10 | Must meet requirement | N / A | Must meet requirement | N / A | Letter of Proposal |
| ## 1.2.3 Pending Litigation | Proposer’s financial position and prospective long-term profitability still sound according to criteria established in 2.3.1 below and assuming that all pending litigation will be resolved against the Proposer. | Must meet requirement by itself or as member to past or existing JV | N / A | Must meet requirement | N / A | Form CON – 2 |
| ## 1.2.4 Litigation History | No consistent history of court/arbitral award decisions against the Proposersince 1st January 2019 | Must meet requirement | Must meet requirement | Must meet requirement | N/A | Form CON – 2 |
| ## 1.2.5 Bank’s SEA and/or SH Disqualification | At the time of Contract Award, not subject to disqualification by the Bank for non-compliance with SEA/ SH obligations | Must meet requirement (including each subcontractor proposed by the Proposer) | N/A | Must meet requirement (including each subcontractor proposed by the Proposer) | N/A | Letter of Proposal, Form CON-3 |

| Factor | # 1.3 Financial Situation |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Sub-Factor | ##### Criteria |  |  |  |  | ##### Documentation Required |
|  | ##### Requirement | ##### Proposer |  |  |  |  |
|  |  | Single Entity | ##### Joint Venture (existing or intended) |  |  |  |
|  |  |  | All members combined | Each member | At least one member |  |
| ## 1.3.1 Historical Financial Performance | Submission of audited balance sheets or if not required by the law of the Proposer’s country, other financial statements acceptable to the Purchaser, for the last three (3) years (2023, 2024, 2025) to demonstrate the current soundness of the Proposers financial position and its prospective long-term profitability. | Must meet requirement | N / A | Must meet requirement | N / A | Form FIN – 1.3.1 with attachments |
| ## 1.3.2 Average Annual Turnover | Minimum average annual turnover of at least USD 3,000,000, calculated as total certified payments received for contracts in progress or completed, within the last three (3) years (2023, 2024, 2025) | Must meet requirement | Must meet requirement | N/A | N/A | Form FIN –1.3.2 |
| ## 1.3.3 Financial Resources | - The Proposer must demonstrate access to, or availability of, financial resources such as liquid assets, unencumbered real assets, lines of credit, and other financial means, other than any contractual advance payments to meet the following cash-flow requirement: Access to USD 750,000.00 | Must meet requirement | Must meet requirement | the Lead member must meet at least 40% of the cash flow requirement, | N/A | Form FIN –1.3.3 |

| Factor | # 1.4 Experience |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Sub-Factor | ##### Criteria |  |  |  |  | ##### Documentation Required |
|  | Requirement | ##### Proposer |  |  |  |  |
|  |  | ##### Single Entity | Joint Venture (existing or intended) |  |  |  |
|  |  |  | All members combined | Each member | At least one member |  |
| ## 1.4.1 General Experience | Experience under Similar contracts in the role of prime supplier, management contractor, JV member, or subcontractor for at least the last five (5) years prior to the applications submission deadline. | Must meet requirement | N / A | Must meet requirement | N / A | Form EXP-1.4.1 |
| ## 1.4.2 Specific Experience | Participation as a prime supplier, management contractor, JV member, sub-contractor, in at least Two (2) Similar contracts within the last five (5) years, each with a value of at least USD 2,500,000, (United States Dollars Two and Half Million) that have been successfully and substantially completed and that are similar to the proposed Supply, Installation and Commissioning. | Must meet requirement | Must meet requirements for all characteristics | N / A | Must meet requirement for one characteristic | Form EXP 1.4.2 |
| ## 1.4.3 Specific Experience | The Proposer must provide at least three (3) Client References for similar successfully completed contracts in public or private sector; at least one reference must be the qualifying contract. Details should include the detailed tasks done, country, year, and Client contact details. Evidence of work done should be supported with Completion Certificates Reference Letters and value of those contracts. | Must meet requirement | Must meet requirements for all characteristics | N / A | Must meet requirement for one characteristic |  |

Section IV - Bidding Forms

Table of Forms

Letter of Bid – Technical Part	84

Technical Part	87

Technical Bid Checklist	88

Functional Guarantees	89

Manufacturer’s Authorization	90

Bidder Information Form	91

Bidder’s JV Members Information Form	93

Sexual Exploitation and Abuse (SEA) and/or Sexual Harassment Performance Declaration	94

Historical Contract Non-Performance, Pending Litigation and Litigation History……………………………95

General Experience………………………………………………………………………………………………………………………..97

Specific Experience………………………………………………………………………………………………………………………..98

Current Contract Commitments / Work in Progress…………………………………………………………………….100

Financial Situation: Historical Financial Performance…………………………………………………………………..101

Average Annual Turnover…………………………………………………………………………………………………………….103

Financial Resources……………………………………………………………………………………………………………………..104

Code of Conduct for Supplier’s Personnel Form…………………………………………………………………………..105

Subcontractor’s Agreement…………………………………………………………………………………………………………114

Form of Bid Security…………………………………………………………………………………………………………………….116

Form of Bid Security (Bid Bond)	117

Form of Bid-Securing Declaration………………………………………………………………………………………………..117

Letter of Bid - Financial Part…………………………………………………………………………………………………………119

Price Schedule Forms…………………………………………………………………………………………………………………..120

Price Schedule: Goods Manufactured Outside the Purchaser’s Country, to be Imported…………….155

Price Schedule: Goods Manufactured Outside the Purchaser’s Country, already imported…………157

Price Schedule: Goods Manufactured in the Purchaser’s Country……………………………………………….159

Price and Completion Schedule - Related Services………………………………………………………………………161

Letter of Bid – Technical Part

INSTRUCTIONS TO BIDDERS: DELETE THIS BOX ONCE YOU HAVE COMPLETED THE DOCUMENT Place this Letter of Bid in the first envelope “TECHNICAL PART”. The Bidder must prepare the Letter of Bid on stationery with its letterhead clearly showing the Bidder’s complete name and business address. Note: All italicized text in black font is to help Bidders in preparing this form and Bidders shall delete it from the final document.

Date of this Bid submission: [insert date (as day, month and year) of Bid submission]

RFB No.: [insert number of Bidding process]

Request for Bid No.: [insert identification]

Alternative No.: [insert identification No if this is a Bid for an alternative]

To: [insert complete name of Purchaser]

We, the undersigned Bidder, hereby submit our Bid, in two parts, namely:

- the Technical Part, and

- the Financial Part.

In submitting our Bid we make the following declarations:

- No reservations: We have examined and have no reservations to the bidding document, including addenda issued in accordance with Instructions to Bidders (ITB 8);

- Eligibility: We meet the eligibility requirements and have no conflict of interest in accordance with ITB 4;

- Bid/Proposal-Securing Declaration: We have not been suspended nor declared ineligible by the Purchaser based on execution of a Bid Securing Declaration  or Proposal Securing Declaration in the Purchaser’s country in accordance with ITB 4.7;

- Sexual Exploitation and Abuse (SEA) and/or Sexual Harassment (SH): [select the appropriate option from (i) to (iii) below and delete the others. In case of JV members and/or subcontractors, indicate the status of disqualification by the Bank of each JV member and/or subcontractor].

We, including any of our subcontractors:

- [have not been subject to disqualification by the Bank for non-compliance with SEA/ SH obligations.]

- [are subject to disqualification by the Bank for non-compliance with SEA/ SH obligations.]

- [had been subject to disqualification by the Bank for non-compliance with SEA/ SH obligations, and were removed from the disqualification list. An arbitral award on the disqualification case has been made in our favor.]

- Conformity: We offer to supply in conformity with the bidding document and in accordance with the Delivery Schedules specified in the Schedule of Requirements the following Goods: [insert a brief description of the Goods and Related Services];

- Bid Validity: Our Bid shall be valid until [insert day, month and year in accordance with ITB 18.1], and it shall remain binding upon us and may be accepted at any time before the expiration of that period;

- Performance Security: If our Bid is accepted, we commit to obtain a performance security in accordance with the bidding document;

- One Bid per Bidder: We are not submitting any other Bid(s) as an individual Bidder, and we are not participating in any other bid(s) as a Joint Venture member or as a subcontractor, and meet the requirements of ITB 4.3, other than Alternative Bids submitted in accordance with ITB 13;

- Suspension and Debarment: We, along with any of our subcontractors, suppliers, consultants, manufacturers, or service providers for any part of the contract, are not subject to, and not controlled by any entity or individual that is subject to, a temporary suspension or a debarment imposed by the World Bank Group or a debarment imposed by the World Bank Group in accordance with the Agreement for Mutual Enforcement of Debarment Decisions between the World Bank and other development banks. Further, we are not ineligible under the Purchaser’s country laws or official regulations or pursuant to a decision of the United Nations Security Council;

- State-owned enterprise or institution: [select the appropriate option and delete the other] [We are not a state-owned enterprise or institution] / [We are a state-owned enterprise or institution but meet the requirements of ITB 4.6];

- Binding Contract: We understand that this Bid, together with your written acceptance thereof included in your Letter of Acceptance, shall constitute a binding contract between us, until a formal contract is prepared and executed;

- Not Bound to Accept: We understand that you are not bound to accept the lowest evaluated cost Bid, the Most Advantageous Bid or any other Bid that you may receive; and

- Fraud and Corruption: We hereby certify that we have taken steps to ensure that no person acting for us, or on our behalf, engages in any type of Fraud and Corruption.

Name of the Bidder: *[insert complete name of Bidder]

Name of the person duly authorized to sign the Bid on behalf of the Bidder: ** [insert complete name of person duly authorized to sign the Bid]

Title of the person signing the Bid: [insert complete title of the person signing the Bid]

Signature of the person named above: [insert signature of person whose name and capacity are shown above]

Date signed [insert date of signing] day of [insert month], [insert year]

*: In the case of the Bid submitted by a Joint Venture specify the name of the Joint Venture as Bidder.

**: Person signing the Bid shall have the power of attorney given by the Bidder. The power of attorney shall be attached with the Bid Schedules.

Technical Part

The technical bid shall include all relevant information required to evaluate the technical Part in accordance with the requirements.

To establish the conformity of the Goods and Related Services to the RFB document, the Bidder shall furnish the documentary evidence that the Goods conform to the technical specifications and standards, including any essential technical and performance characteristics specified in Section VII, Schedule of Requirements. Any required functional guarantees shall also be provided. The attached forms/format may support the Bidder to organize information required to present its technical bid.

The documentary evidence may be in the form of literature, drawings or data, and shall consist of a detailed item by item description of the essential technical and performance characteristics of the Goods and Related Services, demonstrating substantial responsiveness of the Goods and Related Services to the technical specification, and if applicable, a statement of deviations and exceptions to the provisions of the Section VII, Schedule of Requirements.

In the interest of timely bid evaluation and contract award, Bidders are encouraged not to overload the supporting materials with documents that do not directly address the Purchaser’s requirements.

The Bidder shall also furnish a list giving full particulars, including available sources and current prices of spare parts, special tools, etc., necessary for the proper and continuing functioning of the Goods during the period specified in the BDS following commencement of the use of the goods by the Purchaser.

Standards for workmanship, process, material, and equipment, as well as references to brand names or catalogue numbers specified by the Purchaser in the Schedule of Requirements, are intended to be descriptive only and not restrictive. The Bidder may offer other standards of quality, brand names, and/or catalogue numbers, provided that it demonstrates, to the Purchaser’s satisfaction, that the substitutions ensure substantial equivalence or are superior to those specified in the Section VII, Schedule of Requirements.

If the contract has been assessed to present potential or actual cyber security risks, the technical bid must include proposed cyber security risks management plan.

If there are assessed supply chain risks, the technical bid must include proposed supply chain risk management plan.

The Manufacture’s Authorizations shall be included in accordance with ITB BDS 17.2 (a) and the attached Manufacturer’s Authorization form.

Technical Bid Checklist

| Technical. Requirement No. _ | Technical Requirement: [ insert: description of requirement ] |
| --- | --- |
| Bidder’s technical bid/ compliance: |  |
| Bidder’s cross references to supporting information in the Technical Bid: |  |

Functional Guarantees

(to be used as applicable)

The Bidder shall copy in the left column of the table below; the identification of each functional guarantee required in the Specification and in the right column, provides the corresponding value for each functional guarantee of the proposed Goods.

| Required Functional Guarantee | Value of Functional Guarantee of the Goods |
| --- | --- |
| 1. |  |
| 2. |  |
| 3. |  |
| … |  |

Manufacturer’s Authorization

[The Bidder shall require the Manufacturer to fill in this Form in accordance with the instructions indicated. This letter of authorization should be on the letterhead of the Manufacturer and should be signed by a person with the proper authority to sign documents that are binding on the Manufacturer. The Bidder shall include it in its Bid, if so indicated in the BDS.]

Date: [insert date (as day, month and year) of Bid submission]

RFB No.: [insert number of RFB process]

Alternative No.: [insert identification No if this is a Bid for an alternative]

To: [insert complete name of Purchaser]

WHEREAS

We [insert complete name of Manufacturer], who are official manufacturers of [insert type of goods manufactured], having factories at [insert full address of Manufacturer’s factories], do hereby authorize [insert complete name of Bidder] to submit a Bid the purpose of which is to provide the following Goods, manufactured by us [insert name and or brief description of the Goods], and to subsequently negotiate and sign the Contract.

We hereby extend our full guarantee and warranty in accordance with Clause 28 of the General Conditions of Contract, with respect to the Goods offered by the above firm.

We confirm that we do not engage or employ forced labor or persons subject to trafficking or child labor, in accordance with Clause 14 of the General Conditions of Contract.

Signed: [insert signature(s) of authorized representative(s) of the Manufacturer]

Name: [insert complete name(s) of authorized representative(s) of the Manufacturer]

Title: [insert title]

Dated on ____________ day of __________________, _______ [insert date of signing]

Bidder Information Form

[The Bidder shall fill in this Form in accordance with the instructions indicated below. No alterations to its format shall be permitted and no substitutions shall be accepted.]

Date: [insert date (as day, month and year) of Bid submission]

RFB No.: [insert number of Bidding process]

Alternative No.: [insert identification No if this is a Bid for an alternative]

Page ________ of_ ______ pages

1. Bidder’s Name [insert Bidder’s legal name]

2. In case of JV, legal name of each member : [insert legal name of each member in JV]

3. Bidder’s actual or intended country of registration: [insert actual or intended country of registration]

4. Bidder’s year of registration: [insert Bidder’s year of registration]

5. Bidder’s Address in country of registration: [insert Bidder’s legal address in country of registration]

6. Bidder’s Authorized Representative Information Name: [insert Authorized Representative’s name] Address: [insert Authorized Representative’s Address] Telephone/Fax numbers: [insert Authorized Representative’s telephone/fax numbers] Email Address: [insert Authorized Representative’s email address]

7. Attached are copies of original documents of [check the box(es) of the attached original documents] Articles of Incorporation (or equivalent documents of constitution or association), and/or documents of registration of the legal entity named above, in accordance with ITB 4.4. In case of JV, letter of intent to form JV or JV agreement, in accordance with ITB 4.1. In case of state-owned enterprise or institution, in accordance with ITB 4.6 documents establishing: - Legal and financial autonomy - Operation under commercial law - Establishing that the Bidder is not under the supervision of the Purchaser 8. Included are the organizational chart, a list of Board of Directors, and the beneficial ownership. The successful Bidder shall provide additional information on beneficial ownership, using the Beneficial Ownership Disclosure Form.

Bidder’s JV Members Information Form

[The Bidder shall fill in this Form in accordance with the instructions indicated below. The following table shall be filled in for the Bidder and for each member of a Joint Venture]].

Date: [insert date (as day, month and year) of Bid submission]

RFB No.: [insert number of RFB process]

Alternative No.: [insert identification No if this is a Bid for an alternative]

Page ________ of_ ______ pages

1. Bidder’s Name: [insert Bidder’s legal name]

2. Bidder’s JV Member’s name: [insert JV’s Member legal name]

3. Bidder’s JV Member’s country of registration: [insert JV’s Member country of registration]

4. Bidder’s JV Member’s year of registration: [insert JV’s Member year of registration]

5. Bidder’s JV Member’s legal address in country of registration: [insert JV’s Member legal address in country of registration]

6. Bidder’s JV Member’s authorized representative information Name: [insert name of JV’s Member authorized representative] Address: [insert address of JV’s Member authorized representative] Telephone/Fax numbers: [insert telephone/fax numbers of JV’s Member authorized representative] Email Address: [insert email address of JV’s Member authorized representative]

7. Attached are copies of original documents of [check the box(es) of the attached original documents] Articles of Incorporation (or equivalent documents of constitution or association), and/or registration documents of the legal entity named above, in accordance with ITB 4.4. In case of a state-owned enterprise or institution, documents establishing legal and financial autonomy, operation in accordance with commercial law, and that they are not under the supervision of the Purchaser, in accordance with ITB 4.6. 8. Included are the organizational chart, a list of Board of Directors, and the beneficial ownership. The successful Bidder shall provide additional information on beneficial ownership for each JV member using the Beneficial Ownership Disclosure Form.

Sexual Exploitation and Abuse (SEA) and/or Sexual Harassment Performance Declaration

[The following table shall be filled in by the Bidder, each member of a Joint Venture and each subcontractor proposed by the Bidder]

Bidder’s Name: [insert full name]
Date: [insert day, month, year]
Joint Venture Member’s or Subcontractor’s Name: [insert full name]
RFB No. and title: [insert RFB number and title]
Page [insert page number] of [insert total number] pages

SEA and/or SH Declaration in accordance with Section III, Qualification Criteria, and Requirements

We: (a) have not been subject to disqualification by the Bank for non-compliance with SEA/ SH obligations (b) are subject to disqualification by the Bank for non-compliance with SEA/ SH obligations (c) had been subject to disqualification by the Bank for non-compliance with SEA/ SH obligations, and were removed from the disqualification list. An arbitral award on the disqualification case has been made in our favor.

[If (c) above is applicable, attach evidence of an arbitral award reversing the findings on the issues underlying the disqualification.]

NOTE: Please Tick the applicable box and sign the Form.

Historical Contract Non-Performance, Pending Litigation and Litigation History

In case a prequalification process was conducted this form should be used only if the information submitted at the time of prequalification requires updating

Proposer’s Legal Name: [insert Proposer’s Legal Name]

Date:                 [insert date]

JV member Legal Name: [insert JV Member Legal Name]

RFP No.: [insert RFP number]

Page _______ of _______ pages

| Non-Performed Contracts in accordance with Section III, Evaluation and Qualification Criteria |  |  |  |
| --- | --- | --- | --- |
|  Contract non-performance did not occur since 1st January [insert year] specified in Section III, Evaluation and Qualification Criteria, Sub-Factor 1.2.1.  Contract(s) not performed since 1st January [insert year] specified in Section III, Evaluation and Qualification Criteria, requirement 1.2.1 |  |  |  |
| Year | Non- performed portion of contract | Contract Identification | Total Contract Amount (current value, currency, exchange rate and US$ equivalent) |
| [insert year] | [insert amount and percentage] | Contract Identification: [indicate complete contract name/ number, and any other identification] Name of Employer: [insert full name] Address of Employer: [insert street/city/country] Reason(s) for non-performance: [indicate main reason(s)] | [insert amount] |
| Pending Litigation, in accordance with Section III, Evaluation and Qualification Criteria |  |  |  |
|  No pending litigation in accordance with Sub-Factor 1.2.3. |  |  |  |
|  Pending litigation in accordance with Sub-Factor 1.2.3 as indicated below. |  |  |  |

| Year of dispute | Amount in dispute (currency) | Contract Identification | Total Contract Amount (currency), USD Equivalent (exchange rate) |
| --- | --- | --- | --- |
| [specify year] | [specify amount and currency] | Contract Identification: [insert Contract ID] Name of Employer: [insert Name of Employer] Address of Employer: [insert Address of Employer] Matter in dispute: [describe Matter of dispute] Party who initiated the dispute: [specify Initiator of dispute] Status of dispute: [specify Status of dispute] | [specify total contract amount and currency, USD equivalent and exchange rate] |
| [specify year] | [specify amount and currency] | Contract Identification: [insert Contract ID] Name of Employer: [insert Name of Employer] Address of Employer: [insert Address of Employer] Matter in dispute: [describe Matter of dispute] Party who initiated the dispute: [specify Initiator of dispute] Status of dispute: [specify Status of dispute] | [specify total contract amount and currency, USD equivalent and exchange rate] |
| Litigation History in accordance with Section III, Evaluation and Qualification Criteria |  |  |  |
|  No Litigation History in accordance with Sub-Factor 1.2.4.  Litigation History in accordance with Sub-Factor 1.2.4 as indicated below. |  |  |  |
| Year of award | Outcome as percentage of Net Worth | Contract Identification | Total Contract Amount (currency), USD Equivalent (exchange rate) |
| [specify year] | [specify percentage of net worth] | Contract Identification: [insert Contract ID] Name of Employer: [insert Name of Employer] Address of Employer: [insert Address of Employer] | [specify total contract amount and currency, USD equivalent and exchange rate] |

NOTE: Please Tick the applicable box and sign the Form.

Form EXP 1.4.1- General Experience

Proposer’s Legal Name: [insert Proposer’s Legal Name]

Date: [insert Date]

JV Member Legal Name: [insert JV Member Legal Name]

RFP No.: [insert RFP number]

Page _______ of _______ pages

| Starting Month / Year | Ending Month / Year | Years* | Contract Identification | Role of Proposer |
| --- | --- | --- | --- | --- |
| [insert month / year] | [insert month / year] | [insert number of years] | Contract name: [insert Name of Contract] Brief Description of the Information System performed by the Proposer: [describe Information System] Name of Purchaser: [insert Name of Purchaser] Address: [insert Address of Purchaser] | [describe role of Proposer under the contract] |
| [insert month / year] | [insert month / year] | [insert number of years] | Contract name: [insert Name of Contract] Brief Description of the Information System performed by the Proposer: [describe Information System] Name of Purchaser: [insert Name of Purchaser] Address: [insert Address of Purchaser] | [describe role of Proposer under the contract] |
| [insert month / year] | [insert month / year] | [insert number of years] | Contract name: [insert Name of Contract] Brief Description of the Information System performed by the Proposer: [describe Information System] Name of Purchaser: [insert Name of Purchaser] Address: [insert Address of Purchaser] | [describe role of Proposer under the contract] |
| [insert month / year] | [insert month / year] | [insert number of years] | Contract name: [insert Name of Contract] Brief Description of the Information System performed by the Proposer: [describe Information System] Name of Purchaser: [insert Name of Purchaser] Address: [insert Address of Purchaser] | [describe role of Proposer under the contract] |

*List calendar year for years with contracts with at least nine (9) months activity per year starting with the earliest year

Form EXP 1.4.2- Specific Experience

Proposer’s Legal Name: [insert Proposer’s Legal Name]

Date: [insert Date]

JV Member Legal Name: [insert JV Member Legal Name]

RFP No.: [insert RFP number]

Page _______ of _______ pages

| Similar Contract Number: ___ of ___ required. | Information |  |  |
| --- | --- | --- | --- |
| - Contract Identification | - [insert Contract ID] |  |  |
| - Award date - Completion date | - [insert Date of Award - [insert Date of Completion |  |  |
| Role in Contract | Prime Supplier | Management Contractor | Subcontractor |
| - Total contract amount |  |  |  |
| - If member in a JV or subcontractor, specify participation of total contract amount |  |  |  |
| - Purchaser’s Name: | - _____________________________________ |  |  |
| - Address: - Telephone/fax number: - E-mail: | - _____________________________________ - _____________________________________ - _____________________________________ |  |  |

Form EXP 1.4.2 (cont.)- Specific Experience (cont.)

Proposer’s Legal Name: [insert Proposer’s Legal Name]

Date: [insert Date]

JV Member Legal Name: [insert JV Member Legal Name]

RFP No.: [insert RFP number]

Page _______ of _______ pages

| Similar Contract No. [insert specific number] of [insert total number of contracts] required | Information |
| --- | --- |
| Description of the similarity in accordance with Sub-Factor 1.4.2 of Section III: |  |
| Amount | [insert contract amount and currency and USD equivalent and exchange rate] |
| Geographical Scope | [describe geographic scope of the users of the information system] |
| Functional Scope | [describe the functionalities provided by the information system] |
| Methods/Technology | [describe methodologies and technologies used to implement the information system] |
| Key Activities | [describe the key activities of the Proposer under the contract] |

Form CCC- Summary Sheet:  Current Contract Commitments / Work in Progress

[Proposers and each partner to a Joint Venture should provide information on their current commitments on all contracts that have been awarded, or for which a letter of intent or acceptance has been received, or for contracts approaching completion, but for which an unqualified, full completion certificate has yet to be issued.]

Proposer’s Legal Name: [insert Proposer’s Legal Name]

Date: [insert Date]

JV Member Legal Name: [insert JV Member Legal Name]

RFP No.: [insert RFP number]

Page _______ of _______ pages

| Name of contract | Purchaser, contact address/tel./fax | Value of outstanding Information System (current US$ equivalent) | Estimated completion date | Average monthly invoicing over last six months (US$ equivalent/month) |
| --- | --- | --- | --- | --- |
| 1. [insert Name of Contract] | [insert Name of Purchaser, contact address, telephone / fax number] | [insert Total Outstanding Contract Value of the information system in USD equivalent and exchange rate] | [insert Estimated completion date] | [insert Average monthly invoices in USD equivalent and exchange rate] |
| 2. [insert Name of Contract] | [insert Name of Purchaser, contact address, telephone / fax number] | [insert Total Outstanding Contract Value of the information system in USD equivalent and exchange rate] | [insert Estimated completion date] | [insert Average monthly invoices in USD equivalent and exchange rate] |
| 3. [insert Name of Contract] | [insert Name of Purchaser, contact address, telephone / fax number] | [insert Total Outstanding Contract Value of the information system in USD equivalent and exchange rate] | [insert Estimated completion date] | [insert Average monthly invoices in USD equivalent and exchange rate] |
| ... |  |  |  |  |

Form FIN 1.3.1- Financial Situation: Historical Financial Performance

To be completed by the Proposer and, if JV, by each member

Proposer’s Legal Name: [insert Proposer’s Legal Name]

Date: [insert Date]

JV Member Legal Name: [insert JV Member Legal Name]

RFP No.: [insert RFP number]

Page _______ of _______ pages

| Financial information in US$ equivalent | Historic information for previous [insert number] years ##### (US$ equivalent in 000s) |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | Year 1 | Year 2 | Year 3 | Year … | Year n | Avg. | Avg. Ratio |
| Information from Balance Sheet |  |  |  |  |  |  |  |
| Total Assets (TA) |  |  |  |  |  |  |  |
| Total Liabilities (TL) |  |  |  |  |  |  |  |
| Net Worth (NW) |  |  |  |  |  |  |  |
| Current Assets (CA) |  |  |  |  |  |  |  |
| Current Liabilities (CL) |  |  |  |  |  |  |  |
| Information from Income Statement |  |  |  |  |  |  |  |
| Total Revenue (TR) |  |  |  |  |  |  |  |
| Profits Before Taxes (PBT) |  |  |  |  |  |  |  |

Attached are copies of financial statements (balance sheets, including all related notes, and income statements) for the years required above complying with the following conditions:

- Must reflect the financial situation of the Proposer or member to a JV, and not sister or parent companies

- Historic financial statements must be audited by a certified accountant

- Historic financial statements must be complete, including all notes to the financial statements

- Historic financial statements must correspond to accounting periods already completed and audited (no statements for partial periods shall be requested or accepted)

Form FIN 1.3.2- Average Annual Turnover

[To be completed by the Proposer and, if JV, by each member]

Proposer’s Legal Name: [insert Proposer’s Legal Name]

Date: [insert Date]

JV Member Legal Name: [insert JV Member Legal Name]

RFP No.:  [insert RFP number]

Page _______ of _______ pages

| - Annual turnover data (applicable activities only) |  |  |
| --- | --- | --- |
| - Year | - Amount and Currency | - US$ equivalent |
| - [insert year] | - [insert amount and currency] | - [insert amount in USD equivalent and exchange rate] |
| - [insert year] | - [insert amount and currency] | - [insert amount in USD equivalent and exchange rate] |
| - [insert year] | - [insert amount and currency] | - [insert amount in USD equivalent and exchange rate] |
| - [insert year] | - [insert amount and currency] | - [insert amount in USD equivalent and exchange rate] |
| - [insert year] | - [insert amount and currency] | - [insert amount in USD equivalent and exchange rate] |
| - [insert year] | - [insert amount and currency] | - [insert amount in USD equivalent and exchange rate] |
| - *Average Annual Turnover | - [insert amount and currency] | - [insert amount in USD equivalent and exchange rate] |

*Average annual turnover calculated as total certified payments received for work in progress or completed, divided by the number of years specified in Section III, Evaluation and Qualification Criteria, Sub-Factor 1.3.2.

Form FIN 1.3.3- Financial Resources

To be completed by the Proposer and, if JV, by each member

Proposer’s Legal Name:  [insert Proposer’s Legal Name]

Date:  [insert Date]

JV Member Legal Name:  [insert JV Member Legal Name]

RFP No.:  [insert RFP number]

Page _______ of _______ pages

Specify proposed sources of financing, such as liquid assets, unencumbered real assets, lines of credit, and other financial means, net of current commitments, available to meet the total cash flow demands of the subject contract or contracts as indicated in Section III, Evaluation and Qualification Criteria

| Source of financing | Amount (US$ equivalent) |
| --- | --- |
| 1. [describe type and source of available financing] | [insert amount of available financing in USD equivalent and exchange rate] |
| 2. [describe type and source of available financing] | [insert amount of available financing in USD equivalent and exchange rate] |
| 3. [describe type and source of available financing] | [insert amount of available financing in USD equivalent and exchange rate] |
| 4. [describe type and source of available financing] | [insert amount of available financing in USD equivalent and exchange rate] |

Note to the Purchaser:

The following minimum requirements shall not be modified. The Purchaser may add additional requirements to address identified issues, informed by relevant environmental and social assessment.

Delete this Box prior to issuance of the Request for Proposals document.

Code of Conduct for Supplier’s Personnel Form

Note to the Proposer:

The minimum content of the Code of Conduct form as set out by the Purchaser shall not be substantially modified. However, the Proposer may add requirements as appropriate, including to take into account Contract-specific issues/risks.

The Proposer shall initial and submit the Code of Conduct form as part of its proposal.

CODE OF CONDUCT FOR SUPPLIER’S PERSONNEL

We are the Supplier, [enter name of Supplier].  We have signed a contract with [enter name of Purchaser] for [enter description of the Information System]. The Information System will be supplied to and installed at [enter the Project Site/s]. Our contract requires us to implement measures to address environmental and social risks.

This Code of Conduct identifies the behaviour that we require from Supplier’s Personnel employed in the execution of the Contract at the Project Site/s.

Our workplace is an environment where unsafe, offensive, abusive or violent behaviour will not be tolerated and where all persons should feel comfortable raising issues or concerns without fear of retaliation.

REQUIRED CONDUCT

Supplier’s Personnel employed in the execution of the Contract at the Project Site/s shall:

- carry out his/her duties competently and diligently;

- comply with this Code of Conduct and all applicable laws, regulations and other requirements, including requirements to protect the health, safety and well-being of other Supplier’s Personnel and any other person;

- maintain a safe working environment including by:

  - ensuring that workplaces, machinery, equipment and processes under each person’s control are safe and without risk to health;

  - wearing required personal protective equipment;

  - using appropriate measures relating to chemical, physical and biological substances and agents; and

  - following applicable emergency operating procedures.

- report work situations that he/she believes are not safe or healthy and remove himself/herself from a work situation which he/she reasonably believes presents an imminent and serious danger to his/her life or health;

- treat other people with respect, and not discriminate against specific groups such as women, people with disabilities, migrant workers or children;

- not engage in any form of sexual harassment including unwelcome sexual advances, requests for sexual favours, and other verbal or physical conduct of a sexual nature with other Supplier’s or Purchaser’s Personnel;

- not engage in Sexual Exploitation, which means any actual or attempted abuse of position of vulnerability, differential power or trust, for sexual purposes, including, but not limited to, profiting monetarily, socially or politically from the sexual exploitation of another;

- not engage in in Sexual Abuse,  which means the actual or threatened physical intrusion of a sexual nature, whether by force or under unequal or coercive conditions;

- not engage in any form of sexual activity with individuals under the age of 18, except in case of pre-existing marriage;

- complete relevant training courses that will be provided related to the environmental and social aspects of the Contract, including on health and safety matters, and Sexual Exploitation and Abuse, and Sexual Harassment (SH);

- report violations of this Code of Conduct; and

- not retaliate against any person who reports violations of this Code of Conduct, whether to us or the Purchaser, or who makes use of the grievance mechanism for Supplier’s Personnel or the project’s Grievance Redress Mechanism.

RAISING CONCERNS

If any person observes behaviour that he/she believes may represent a violation of this Code of Conduct, or that otherwise concerns him/her, he/she should raise the issue promptly. This can be done in either of the following ways:

- Contact [enter name of the Supplier’s Social Expert with relevant experience in handling sexual exploitation, sexual abuse and sexual harassment cases, or if such person is not required under the Contract, another individual designated by the Supplier to handle these matters] in writing at this address [   ] or by telephone at [   ] or in person at [   ]; or

- Call [  ]  to reach the Supplier’s hotline (if any) and leave a message.

The person’s identity will be kept confidential, unless reporting of allegations is mandated by the country law. Anonymous complaints or allegations may also be submitted and will be given all due and appropriate consideration. We take seriously all reports of possible misconduct and will investigate and take appropriate action. We will provide warm referrals to service providers that may help support the person who experienced the alleged incident, as appropriate.

There will be no retaliation against any person who raises a concern in good faith about any behaviour prohibited by this Code of Conduct.  Such retaliation would be a violation of this Code of Conduct.

CONSEQUENCES OF VIOLATING THE CODE OF CONDUCT

Any violation of this Code of Conduct by the Supplier’s Personnel may result in serious consequences, up to and including termination and possible referral to legal authorities.

FOR SUPPLIER’s PERSONNEL:

I have received a copy of this Code of Conduct written in a language that I comprehend.  I understand that if I have any questions about this Code of Conduct, I can contact [enter name of Supplier’s contact person(s) with relevant experience] requesting an explanation.

Name of Supplier’s Personnel: [insert name]

Signature: __________________________________________________________

Date: (day month year): [insert date]

Countersignature of authorized representative of the Supplier:

Signature: ________________________________________________________

Date: (day month year): [insert date]

ATTACHMENT 1: Behaviours constituting SEA and behaviours constituting SH

ATTACHMENT 1 TO THE CODE OF CONDUCT FORM

BEHAVIORS CONSTITUTING SEXUAL EXPLOITATION AND ABUSE (SEA) AND BEHAVIORS CONSTITUTING SEXUAL HARASSMENT (SH)

The following non-exhaustive list is intended to illustrate types of prohibited behaviours.

- Examples of sexual exploitation and abuse include, but are not limited to:

- A Supplier’s Personnel tells a member of the community that he/she can get them jobs related to the project site in exchange for sex.

- A Supplier’s Personnel rapes, or otherwise sexually assaults a member of the community.

- A Supplier’s Personnel denies a person access to the Project Site/s unless he/she performs a sexual favour.

- A Supplier’s Personnel tells a person applying for employment under the Contract that he/she will only hire him/her if he/she has sex with him/her.

- Examples of sexual harassment in a work context

- A Supplier’s Personnel comment on the appearance of another Supplier’s Personnel (either positive or negative) and sexual desirability.

- When a Supplier’s Personnel complains about comments made by another Supplier’s Personnel on his/her appearance, the other Supplier’s Personnel comment that he/she is “asking for it” because of how he/she dresses.

- Unwelcome touching of a Supplier’s Personnel or Purchaser’s Personnel by another Supplier’s Personnel.

- A Supplier’s Personnel tells another Supplier’s Personnel that he/she will get him/her a salary raise, or promotion if he/she sends him/her naked photographs of himself/herself.

Subcontractor’s Agreement

Note: 	This agreement should be written on the letterhead of the Subcontractor and be signed by a person with the proper authority to sign documents that are binding on the Subcontractor.

Invitation for Proposals Title and No.:  [Purchaser insert:  RFP Title and Number]

To:  [Purchaser insert:  Purchaser’s Officer to receive the Subcontractor’s Agreement]

WHEREAS [ insert: Name of Subcontractor ], having head offices at [ insert: address of Subcontractor ], have been informed by [ insert:  name of Proposer or Joint Venture ] located at [ insert:  address of Proposer or Joint Venture ] (hereinafter, the “Proposer”) that it will submit a proposal in which [ insert: Name of Subcontractor ] will provide  [ insert: items of supply or services provided by the Subcontractor ].   We hereby commit to provide the above named items, in the instance that the Proposal is awarded the Contract.

Name [insert: Name of Officer] in the capacity of [insert: Title of Officer]

Signed  ______________________________

Duly authorized to sign the authorization for and on behalf of: [insert: Name of Subcontractor]

Dated this [ insert: ordinal ] day of  [ insert: month ], [ insert: year ].

[add Corporate Seal (where appropriate)]

List of Proposed Subcontractors

|  | Item | Proposed Subcontractor | Place of Registration & Qualifications |
| --- | --- | --- | --- |

Form of Bid Security

(Bank Guarantee)

(NOT APPLICABLE)

[The bank shall fill in this Bank Guarantee Form in accordance with the instructions indicated.]

[Guarantor letterhead or SWIFT identifier code]

Beneficiary: [Purchaser to insert its name and address]

RFB No.: [Purchaser to insert reference number for the Request for Bids]

Alternative No.: [Insert identification No if this is a Bid for an alternative]

Date: [Insert date of issue]

BID GUARANTEE No.: [Insert guarantee reference number]

Guarantor: [Insert name and address of place of issue, unless indicated in the letterhead]

We have been informed that ______ [insert name of the Bidder, which in the case of a joint venture shall be the name of the joint venture (whether legally constituted or prospective) or the names of all members thereof] (hereinafter called "the Applicant") has submitted or will submit to the Beneficiary its Bid (hereinafter called "the Bid") for the execution of ________________ under Request for Bids No. __________ _ (“the RFB”).

Furthermore, we understand that, according to the Beneficiary’s conditions, Bids must be supported by a Bid guarantee.

At the request of the Applicant, we, as Guarantor, hereby irrevocably undertake to pay the Beneficiary any sum or sums not exceeding in total an amount of ___________ (____________) upon receipt by us of the Beneficiary’s complying demand, supported by the Beneficiary’s statement, whether in the demand itself or a separate signed document accompanying or identifying the demand, stating that either the Applicant:

(a) 	has withdrawn its Bid prior to the Bid validity expiry date set forth in the Applicant’s Letter of Bid, or any extended date provided by the Applicant; or

(b) 	having been notified of the acceptance of its Bid by the Beneficiary prior to the expiry date of the Bid validity or any extension thereof provided by the Applicant has failed to: (i) sign the contract agreement, or (ii) furnish the performance security, in accordance with the Instructions to Bidders (“ITB”) of the Beneficiary’s bidding document.

This guarantee will expire: (a) if the Applicant is the successful Bidder, upon our receipt of copies of the Contract agreement signed by the Applicant and the performance security issued to the Beneficiary in relation to such Contract agreement; or (b) if the Applicant is not the successful Bidder, upon the earlier of (i) our receipt of a copy of the Beneficiary’s notification to the Applicant of the results of the Bidding process; or (ii) twenty-eight days after the expiry date of the Bid validity.

Consequently, any demand for payment under this guarantee must be received by us at the office indicated above on or before that date.

This guarantee is subject to the Uniform Rules for Demand Guarantees (URDG) 2010 Revision, ICC Publication No. 758.

_____________________________

[Signature(s)]

Note: All italicized text is for use in preparing this form and shall be deleted from the final product.

Form of Bid Security (Bid Bond)

(NOT APPLICABLE)

[The Surety shall fill in this Bid Bond Form in accordance with the instructions indicated.]

BOND NO. ______________________

BY THIS BOND [name of Bidder] as Principal (hereinafter called “the Principal”), and [name, legal title, and address of surety], authorized to transact business in [name of country of Purchaser], as Surety (hereinafter called “the Surety”), are held and firmly bound unto [name of Purchaser] as Obligee (hereinafter called “the Purchaser”) in the sum of [amount of Bond] [amount in words], for the payment of which sum, well and truly to be made, we, the said Principal and Surety, bind ourselves, our successors and assigns, jointly and severally, firmly by these presents.

WHEREAS the Principal has submitted or will submit a written Bid to the Purchaser dated the ___ day of ______, 20__, for the supply of [name of Contract] (hereinafter called the “Bid”).

NOW, THEREFORE, THE CONDITION OF THIS OBLIGATION is such that if the Principal:

- withdraws its Bid prior to the Bid validity expiry date set forth in the Principal’s Letter of Bid, or any extended date provided by the Principal; or

- having been notified of the acceptance of its Bid by the Purchaser prior to the expiry date of the Bid validity or any extension thereto provided by the Applicant has failed to: (i) execute the Contract agreement; or (ii) furnish the Performance Security, in accordance with the Instructions to Bidders (“ITB”) of the Purchaser’s bidding document.

then the Surety undertakes to immediately pay to the Purchaser up to the above amount upon receipt of the Purchaser’s first written demand, without the Purchaser having to substantiate its demand, provided that in its demand the Purchaser shall state that the demand arises from the occurrence of any of the above events, specifying which event(s) has occurred.

The Surety hereby agrees that its obligation will remain in full force and effect up to and including the date 28 days after the date of expiry of the Bid validity set forth in the Principal’s Letter of Bid or any extension thereto provided by the Principal.

IN TESTIMONY WHEREOF, the Principal and the Surety have caused these presents to be executed in their respective names this ____ day of ____________ 20__.

Principal: _______________________	Surety: _____________________________

Apply Corporate Seal (where appropriate)

_______________________________	____________________________________ 
(Signature)	(Signature) 
(Printed name and title)	(Printed name and title)

Form of Bid-Securing Declaration

[The Bidder shall fill in this Form in accordance with the instructions indicated.]

Date: [date (as day, month and year)]

RFB No.: [number of RFB process]

Alternative No.: [insert identification No if this is a Bid for an alternative]

To: [complete name of Purchaser]

We, the undersigned, declare that:

We understand that, according to your conditions, Bids must be supported by a Bid-Securing Declaration.

We accept that we will automatically be suspended from being eligible for Bidding or submitting proposals in any contract with the Purchaser for the period of time specified in Section II – Bid Data Sheet if we are in breach of our obligation(s) under the Bid conditions, because we:

(a) 	have withdrawn our Bid prior to the expiry date of the Bid validity specified in the Letter of Bid or any extended date provided by us; or

(b) 	having been notified of the acceptance of our Bid by the Purchaser prior to the expiry date of the Bid validity in the Letter of Bid or any extended date provided by us, (i) fail or refuse to sign the Contract; or (ii) fail or refuse to furnish the Performance Security, if required, in accordance with the ITB.

We understand this Bid Securing Declaration shall expire if we are not the successful Bidder, upon the earlier of (i) our receipt of your notification to us of the name of the successful Bidder; or (ii) twenty-eight days after the expiry date of the Bid validity.

Name of the Bidder*

Name of the person duly authorized to sign the Bid on behalf of the Bidder**	_______

Title of the person signing the Bid	______________________

Signature of the person named above	______________________

Date signed ________________________________ day of ___________________, _____

*: In the case of the Bid submitted by joint venture specify the name of the Joint Venture as Bidder

**: Person signing the Bid shall have the power of attorney given by the Bidder attached to the Bid

[Note: In case of a Joint Venture, the Bid-Securing Declaration must be in the name of all members to the Joint Venture that submits the Bid.]

Letter of Bid - Financial Part

INSTRUCTIONS TO BIDDERS: DELETE THIS BOX ONCE YOU HAVE COMPLETED THE DOCUMENT Place this Letter of Bid - Financial Part in the second envelope marked “FINANCIAL PART”. The Bidder must prepare the Letter of Bid - Financial Part on stationery with its letterhead clearly showing the Bidder’s complete name and business address. Note: All italicized text is to help Bidders in preparing this form.

Date of this Bid submission: [insert date (as day, month and year) of Bid submission]

RFB No.: [insert number of bidding process]

Request for Bid No.: [insert identification]

Alternative No.: [insert identification No if this is a Bid for an alternative]

To: [insert complete name of Purchaser]

We, the undersigned Bidder, hereby submit the second part of our Bid, the Financial Part

In submitting our Financial Part we make the following additional declarations:

- Bid Validity: Our Bid shall be valid until [insert day, month and year in accordance with ITB 18.1], and it shall remain binding upon us and may be accepted at any time before the expiration of that period;

- Total Price: The total price of our Bid, excluding any discounts offered in item (c) below is:

In case of only one lot, the total price of the Bid is [insert the total price of the bid in words and figures, indicating the various amounts and the respective currencies];

In case of multiple lots, the total price of each lot is [insert the total price of each lot in words and figures, indicating the various amounts and the respective currencies];

In case of multiple lots, total price of all lots (sum of all lots) [insert the total price of all lots in words and figures, indicating the various amounts and the respective currencies];

- Discounts: The discounts offered and the methodology for their application are:

(i) The discounts offered are: [Specify in detail each discount offered]

(ii) The exact method of calculations to determine the net price after application of discounts is shown below: [Specify in detail the method that shall be used to apply the discounts];

- Commissions, gratuities and fees: We have paid, or will pay the following commissions, gratuities, or fees with respect to the bidding process or execution of the Contract: [insert complete name of each Recipient, its full address, the reason for which each commission or gratuity was paid and the amount and currency of each such commission or gratuity].

| Name of Recipient | Address | Reason | Amount |
| --- | --- | --- | --- |

(If none has been paid or is to be paid, indicate “none.”)

- Binding Contract: We understand that this Bid, together with your written acceptance thereof included in your Letter of Acceptance, shall constitute a binding contract between us, until a formal contract is prepared and executed.

Name of the Bidder:*[insert complete name of the Bidder]

Name of the person duly authorized to sign the Bid on behalf of the Bidder: ** [insert complete name of person duly authorized to sign the Bid]

Title of the person signing the Bid: [insert complete title of the person signing the Bid]

Signature of the person named above: [insert signature of person whose name and capacity are shown above]

Date signed [insert date of signing] day of [insert month], [insert year]

*: In the case of the Bid submitted by a Joint Venture specify the name of the Joint Venture as Bidder.

**: Person signing the Bid shall have the power of attorney given by the Bidder. The power of attorney shall be attached with the Bid Schedules.

Price Schedule Forms

[The Bidder shall fill in these Price Schedule Forms in accordance with the instructions indicated. The list of line items in column 1 of the Price Schedules shall coincide with the List of Goods and Related Services specified by the Purchaser in the Schedule of Requirements.]

| Price Schedule: Goods Manufactured Outside the Purchaser’s Country, to be Imported |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | (Group C Bids, goods to be imported) Currencies in accordance with ITB 15 |  |  | Date:_________________________ RFB No: _____________________ Alternative No: ________________ Page N ______ of ______ |  |
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 |
| Line Item N | Description of Goods | Country of Origin | Delivery Date as defined by Incoterms | Quantity and physical unit | Unit price DDP – SoBS National Reference Laboratory (NRL), Mogadishu, Somalia in accordance with ITB 14.8(b)(i) | DDP Price per line item (Col. 5x6) | Price per line item for inland transportation and other services required in the Purchaser’s Country to convey the Goods to their final destination specified in BDS | Total Price per Line item (Col. 7+8) |
| A. Analytical Chemistry Equipment |  |  |  |  |  |  |  |  |
| 1 | LC-MS/MS System (Triple Quadrupole) |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| 2 | GC-MS/MS System (Triple Quadrupole) |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| 3 | HPLC System (Quaternary Gradient, DAD + FLD) |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| 4 | GC-MS System (Single Quadrupole + FID/ECD) |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| 5 | AAS – Flame + Graphite Furnace (Zeeman) |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| 6 | FTIR Spectrometer (Benchtop ATR) |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| 7 | UV/Vis Spectrophotometer (Double-Beam, Scanning) |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| 8 | NIR Analyzer (Rapid Compositional, Benchtop) |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| 9 | Ion Chromatograph (Suppressed Conductivity) |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| 10 | Karl Fischer Titrator (Coulometric) |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| 11 | HPLC Column – C18 Reverse Phase (150×4.6 mm, 5 µm) |  | 20 weeks from contract effectiveness | 6 No. |  |  |  |  |
| 12 | HPLC Column – C8 (150×4.6 mm, 5 µm) |  | 20 weeks from contract effectiveness | 4 No. |  |  |  |  |
| 13 | HPLC Column – HILIC (100×2.1 mm, 1.8 µm) |  | 20 weeks from contract effectiveness | 4 No. |  |  |  |  |
| 14 | HPLC Column – Phenyl-Hexyl (150×4.6 mm, 5 µm) |  | 20 weeks from contract effectiveness | 4 No. |  |  |  |  |
| 15 | HPLC Guard Columns (C18, 12.5×4.6 mm) – Pack of 3 |  | 20 weeks from contract effectiveness | 12 No. |  |  |  |  |
| 16 | GC Capillary Column – DB-5 equivalent (30 m × 0.25 mm × 0.25 µm) |  | 20 weeks from contract effectiveness | 4 No. |  |  |  |  |
| 17 | GC Capillary Column – Mid-polarity (30 m × 0.25 mm × 0.25 µm) |  | 20 weeks from contract effectiveness | 2 No. |  |  |  |  |
| 18 | GC Capillary Column – Wax-phase (30 m × 0.25 mm × 0.25 µm) |  | 20 weeks from contract effectiveness | 2 No. |  |  |  |  |
| 19 | SPE Cartridges – C18 (200 mg / 3 mL), Box 50 |  | 20 weeks from contract effectiveness | 20 No. |  |  |  |  |
| 20 | SPE Cartridges – Mixed-mode Dispersive (QuEChERS), Box 50 |  | 20 weeks from contract effectiveness | 20 No. |  |  |  |  |
| 21 | Immunoaffinity Columns – Aflatoxin (25/box) |  | 20 weeks from contract effectiveness | 8 No. |  |  |  |  |
| 22 | Immunoaffinity Columns – Ochratoxin A (25/box) |  | 20 weeks from contract effectiveness | 6 No. |  |  |  |  |
| 23 | Immunoaffinity Columns – Fumonisin B1/B2 (25/box) |  | 20 weeks from contract effectiveness | 6 No. |  |  |  |  |
| B. Microbiology Equipment |  |  |  |  |  |  |  |  |
| 24 | Biosafety Cabinet – Class II A2 (1.2 m) |  | 20 weeks from contract effectiveness | 4 No. |  |  |  |  |
| 25 | Autoclave / Steam Sterilizer (100 L, Pre-Vacuum) |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| 26 | Anaerobic Incubator / Workstation |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| 27 | Incubator – Temperature Controlled (100–400 L) |  | 20 weeks from contract effectiveness | 4 No. |  |  |  |  |
| 28 | Real-Time PCR System (96-well, 4–6 channel) |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| 29 | Deep Freezer –80°C (Ultra-Low, 400 L) |  | 20 weeks from contract effectiveness | 2 No. |  |  |  |  |
| 30 | Refrigerator (2–8°C, 250 L, Lab-Grade) |  | 20 weeks from contract effectiveness | 4 No. |  |  |  |  |
| 31 | Automated Microplate Reader (ELISA) |  | 20 weeks from contract effectiveness | 3 No. |  |  |  |  |
| 32 | Colony Counter (Digital, Image Analysis) |  | 20 weeks from contract effectiveness | 5 No. |  |  |  |  |
| 33 | Fluorescence Microscope (Upright) |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| 34 | Impedance / ATP Microbial Analyzer |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| 35 | Culture Media – Annual Supply (Selective + Non-selective, 14 types) |  | 20 weeks from contract effectiveness | 150 No. |  |  |  |  |
| 36 | Reference Cultures – Freeze-dried (8 species, certified) |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| C. Sample Preparation and General Laboratory Equipment |  |  |  |  |  |  |  |  |
| 37 | Analytical Balance (High-Precision, 220 g / 0.01 mg) |  | 20 weeks from contract effectiveness | 2 No. |  |  |  |  |
| 38 | Precision Balance (50 kg / 1 g) |  | 20 weeks from contract effectiveness | 2 No. |  |  |  |  |
| 39 | pH Meter (Bench-top, Multi-parameter, IP54) |  | 20 weeks from contract effectiveness | 4 No. |  |  |  |  |
| 40 | Moisture Analyzer (Halogen/IR) |  | 20 weeks from contract effectiveness | 3 No. |  |  |  |  |
| 41 | Refractometer (Digital, Peltier-controlled) |  | 20 weeks from contract effectiveness | 2 No. |  |  |  |  |
| 42 | Homogenizer / Stomacher – 400 mL (Paddle Blender) |  | 20 weeks from contract effectiveness | 2 No. |  |  |  |  |
| 43 | Stomacher – High-Volume (3500 mL) |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| 44 | High-Speed Centrifuge – Refrigerated (0–25,000 rpm) |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| 45 | Microcentrifuge (Benchtop, 0–16,000 rpm) |  | 20 weeks from contract effectiveness | 2 No. |  |  |  |  |
| 46 | Water Bath / Heating Bath (5–99°C) |  | 20 weeks from contract effectiveness | 3 No. |  |  |  |  |
| 47 | Dry Bath / Block Heater (Ambient+5 to 130°C) |  | 20 weeks from contract effectiveness | 3 No. |  |  |  |  |
| 48 | Orbital Shaker / Incubator Shaker |  | 20 weeks from contract effectiveness | 2 No. |  |  |  |  |
| 49 | Vortex Mixer (Variable 100–3200 rpm) |  | 20 weeks from contract effectiveness | 4 No. |  |  |  |  |
| 50 | Magnetic Stirrer / Hot Plate (Combined) |  | 20 weeks from contract effectiveness | 6 No. |  |  |  |  |
| 51 | Vacuum Oven (100 L, Ambient–250°C) |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| 52 | Muffle Furnace – Ash Determination (1200°C max) |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| 53 | Drying Oven – Forced Air (Ambient to 300°C) |  | 20 weeks from contract effectiveness | 3 No. |  |  |  |  |
| 54 | Ultrasonic Cleaner (20 L, 40 kHz) |  | 20 weeks from contract effectiveness | 2 No. |  |  |  |  |
| 55 | Water Purification System (Type I + Type II) |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| 56 | Turbidimeter / Nephelometer (0–4000 NTU) |  | 20 weeks from contract effectiveness | 3 No. |  |  |  |  |
| 57 | Dissolved Oxygen / Multi-parameter Meter |  | 20 weeks from contract effectiveness | 2 No. |  |  |  |  |
| 58 | Water Activity Meter (0–1.00 aw) |  | 20 weeks from contract effectiveness | 2 No. |  |  |  |  |
| 59 | Texture Analyzer (0.01–250 N) |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| 60 | Colorimeter (Meat/Food Colour) |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| 61 | Pipette Set – Single Channel (6 volume ranges) |  | 20 weeks from contract effectiveness | 6 No. |  |  |  |  |
| 62 | Multi-Channel Pipette (8 & 12-channel, 50–300 µL) |  | 20 weeks from contract effectiveness | 4 No. |  |  |  |  |
| D. Infrastructure and Support Equipment |  |  |  |  |  |  |  |  |
| 63 | UPS System – Online, 10 kVA (for LC-MS/MS, HPLC, AAS) |  | 20 weeks from contract effectiveness | 4 No. |  |  |  |  |
| 64 | UPS System – Online, 3 kVA (for PCR, Freezers, Incubators) |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| 65 | Power Line Conditioner / Voltage Stabilizer (5 kVA) |  | 20 weeks from contract effectiveness | 6 No. |  |  |  |  |
| 66 | HVAC – Dedicated Lab Precision Air Handling Unit |  | 20 weeks from contract effectiveness | 2 No. |  |  |  |  |
| 67 | Fume Hood – Ducted Chemical (1.2–1.5 m) |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| 68 | Sample Reception Area Setup |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| 69 | Cold Storage Tracking System (Integrated with LIMS) |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| 70 | Sample Storage Racks – Traceable (Cryogenic + Ambient) |  | 20 weeks from contract effectiveness | 20 No. |  |  |  |  |
| 71 | Refrigerator -20°C (Intermediate Sample Storage, 200 L) |  | 20 weeks from contract effectiveness | 2 No. |  |  |  |  |
| 72 | Cold Chain Maintenance (Portable Coolers, PCM Ice Packs) |  | 20 weeks from contract effectiveness | 10 No. |  |  |  |  |
| 73 | Data Logger – Temperature/Humidity (USB, Alarm) |  | 20 weeks from contract effectiveness | 10 No. |  |  |  |  |
| 74 | Emergency Shower & Eyewash Station – Plumbed (ANSI Z358.1) |  | 20 weeks from contract effectiveness | 3 No. |  |  |  |  |
| 75 | Portable Eyewash Station (Self-contained, gravity-fed) |  | 20 weeks from contract effectiveness | 4 No. |  |  |  |  |
| 76 | Chemical Spill Kit (Acid + Base + Solvent) |  | 20 weeks from contract effectiveness | 8 No. |  |  |  |  |
| 77 | Biohazard Waste Containers (Autoclavable Bags + Sharps) |  | 20 weeks from contract effectiveness | 20 No. |  |  |  |  |
| 78 | Chemical Waste Storage Cabinet (Acid-Resistant, 90 L) |  | 20 weeks from contract effectiveness | 4 No. |  |  |  |  |
| 79 | CO₂ Fire Extinguisher (5 kg, EN 3) |  | 20 weeks from contract effectiveness | 10 No. |  |  |  |  |
| 80 | Photoelectric Smoke Detector (IP44, Lab-grade) |  | 20 weeks from contract effectiveness | 8 No. |  |  |  |  |
| 81 | First Aid Kit (Laboratory Grade, 50-person) |  | 20 weeks from contract effectiveness | 4 No. |  |  |  |  |
| E. Enabling Systems (IT, LIMS and Calibration Support Equipment) |  |  |  |  |  |  |  |  |
| 82 | Barcode Label Printer (Lab-grade, 300 dpi, Thermal) |  | 20 weeks from contract effectiveness | 4 No. |  |  |  |  |
| 83 | Barcode Scanner (USB + Wireless, 1D/2D) |  | 20 weeks from contract effectiveness | 8 No. |  |  |  |  |
| 84 | Role-Based Access Control System (LIMS + Instrument Access) |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| 85 | Environmental Monitoring Sensors (Wireless, Temp/RH) |  | 20 weeks from contract effectiveness | 20 No. |  |  |  |  |
| 86 | Certified Weight Set – OIML Class E2/F1 (1 mg to 2 kg) |  | 20 weeks from contract effectiveness | 3 No. |  |  |  |  |
| 87 | Pipette Calibration System – Gravimetric (ISO 8655) |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| 88 | Temperature Calibration – Dry Block Calibrator + Reference Thermometer |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| 89 | Calibration Thermometer – Handheld, 4-channel, Pt100 |  | 20 weeks from contract effectiveness | 4 No. |  |  |  |  |
| 90 | Laboratory Server / NAS Backup System |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| 91 | Network Infrastructure (Managed Switch + Lab Wi-Fi) |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| 92 | Analyst Workstation – Instrument-Dedicated (Desktop) |  | 20 weeks from contract effectiveness | 4 No. |  |  |  |  |
| 93 | Lab Analyst PC – LIMS & Reporting (Desktop) |  | 20 weeks from contract effectiveness | 6 No. |  |  |  |  |
| 94 | Lab Manager / QA Officer Laptop |  | 20 weeks from contract effectiveness | 2 No. |  |  |  |  |
| 95 | Network Laser Printer – A4/A3 (Mono + Colour, Duplex) |  | 20 weeks from contract effectiveness | 2 No. |  |  |  |  |
| 96 | UPS – Line-Interactive (1 kVA) for Workstations & PCs |  | 20 weeks from contract effectiveness | 12 No. |  |  |  |  |
| 97 | External Hard Drive – Encrypted Backup (4 TB, USB 3.0) |  | 20 weeks from contract effectiveness | 6 No. |  |  |  |  |
| F. Livestock and Feed Testing Equipment |  |  |  |  |  |  |  |  |
| 98 | Milk Analyzer – Multi-parameter (Fat/Protein/Lactose/SNF/Conductivity) |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| 99 | Hematology Analyzer – Automated CBC (Livestock Blood Panels) |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| 100 | Veterinary Biochemistry Analyzer (Photometric, Automated) |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| 101 | Portable Milk Tester – Field Grade |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| 102 | Gerber Centrifuge (12-position, ISO 2446) |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| 103 | Portable Livestock Disease Rapid Test Reader |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| 104 | Stereomicroscope (Feed Adulteration Examination) |  | 20 weeks from contract effectiveness | 2 No. |  |  |  |  |
| 105 | Feed Particle Sieve Analyser (Vibrating Sieve Tower) |  | 20 weeks from contract effectiveness | 1 No. |  |  |  |  |
| 106 | ELISA Kit – Brucellosis (B. abortus + B. melitensis) |  | 20 weeks from contract effectiveness | 5 No. |  |  |  |  |
| 107 | ELISA Kit – Foot & Mouth Disease (FMD, Serotypes O/A/SAT) |  | 20 weeks from contract effectiveness | 5 No. |  |  |  |  |
| 108 | ELISA Kit – African Swine Fever (ASFV Antigen + Antibody) |  | 20 weeks from contract effectiveness | 10 No. |  |  |  |  |
| 109 | ELISA Kit – Peste des Petits Ruminants (PPR, cELISA) |  | 20 weeks from contract effectiveness | 10 No. |  |  |  |  |
| 110 | ELISA Kit – Rift Valley Fever (RVF, IgM/IgG) |  | 20 weeks from contract effectiveness | 5 No. |  |  |  |  |
| 111 | ELISA Kit – Bovine Tuberculosis (IFN-gamma, 24-well) |  | 20 weeks from contract effectiveness | 10 No. |  |  |  |  |
| 112 | ELISA Kit – Avian Influenza H5/H7 (Competitive ELISA) |  | 20 weeks from contract effectiveness | 10 No. |  |  |  |  |
| 113 | Lateral Flow Strip – Beta-Lactam Antibiotics (Milk + Tissue) |  | 20 weeks from contract effectiveness | 4 No. |  |  |  |  |
| 114 | Lateral Flow Strip – Tetracyclines (4-in-1 card) |  | 20 weeks from contract effectiveness | 4 No. |  |  |  |  |
| 115 | Lateral Flow Strip – Chloramphenicol (Banned Substance) |  | 20 weeks from contract effectiveness | 2 No. |  |  |  |  |
| 116 | Lateral Flow Strip – Streptomycin / Aminoglycosides |  | 20 weeks from contract effectiveness | 2 No. |  |  |  |  |
| 117 | Lateral Flow Strip – Clenbuterol (Beta-agonist Growth Promoter) |  | 20 weeks from contract effectiveness | 4 No. |  |  |  |  |
| 118 | Lateral Flow Strip – Aflatoxin M1 in Milk |  | 20 weeks from contract effectiveness | 4 No. |  |  |  |  |
| 119 | Blood Collection Tubes – EDTA (10 mL, 1000/box) |  | 20 weeks from contract effectiveness | 5 No. |  |  |  |  |
| 120 | Blood Collection Tubes – Plain/SST (10 mL, 1000/box) |  | 20 weeks from contract effectiveness | 5 No. |  |  |  |  |
| 121 | Vacutainer Needles + Holders (21G × 1.5", 100/box) |  | 20 weeks from contract effectiveness | 30 No. |  |  |  |  |
| 122 | Animal Ear Tag Labels + Barcode Stickers (1000/roll) |  | 20 weeks from contract effectiveness | 10 No. |  |  |  |  |
| 123 | Feed Sampling Probe (Stainless Steel, 1.5 m) |  | 20 weeks from contract effectiveness | 5 No. |  |  |  |  |
| 124 | Feed CRM – Proximate Analysis (Protein/Fat/Moisture/Ash/Fibre) |  | 20 weeks from contract effectiveness | 10 No. |  |  |  |  |
| 125 | Milk CRM – Compositional Analysis (Fat/Protein/Lactose) |  | 20 weeks from contract effectiveness | 6 No. |  |  |  |  |
| G. Consumables and Reagents (First-Year Supply) |  |  |  |  |  |  |  |  |
| 126 | Acetonitrile – HPLC Grade (2.5 L) |  | 20 weeks from contract effectiveness | 20 No. |  |  |  |  |
| 127 | Methanol – HPLC Grade (2.5 L) |  | 20 weeks from contract effectiveness | 20 No. |  |  |  |  |
| 128 | Hexane – HPLC Grade (2.5 L) |  | 20 weeks from contract effectiveness | 15 No. |  |  |  |  |
| 129 | Acetone – HPLC Grade (2.5 L) |  | 20 weeks from contract effectiveness | 15 No. |  |  |  |  |
| 130 | Dichloromethane – HPLC Grade (2.5 L) |  | 20 weeks from contract effectiveness | 10 No. |  |  |  |  |
| 131 | Ethanol – AR Grade (2.5 L) |  | 20 weeks from contract effectiveness | 15 No. |  |  |  |  |
| 132 | Acetic Acid – Glacial AR (2.5 L) |  | 20 weeks from contract effectiveness | 10 No. |  |  |  |  |
| 133 | Nitric Acid – Trace Metal Grade (500 mL) |  | 20 weeks from contract effectiveness | 8 No. |  |  |  |  |
| 134 | Hydrochloric Acid – AR Grade (2.5 L) |  | 20 weeks from contract effectiveness | 10 No. |  |  |  |  |
| 135 | Sulfuric Acid – AR Grade (2.5 L) |  | 20 weeks from contract effectiveness | 10 No. |  |  |  |  |
| 136 | Sodium Hydroxide – AR Pellets (1 kg) |  | 20 weeks from contract effectiveness | 10 No. |  |  |  |  |
| 137 | Heavy Metal Standards – 1000 mg/L (8 elements) |  | 20 weeks from contract effectiveness | 24 No. |  |  |  |  |
| 138 | Pesticide Reference Standard Mix (100+ compounds) |  | 20 weeks from contract effectiveness | 6 No. |  |  |  |  |
| 139 | Veterinary Drug Standard Mix (50–100 compounds) |  | 20 weeks from contract effectiveness | 6 No. |  |  |  |  |
| 140 | Mycotoxin Standards (AFB1/B2/G1/G2, OTA, Fumonisin B1/B2) |  | 20 weeks from contract effectiveness | 12 No. |  |  |  |  |
| 141 | Certified Reference Material – Meat/Fish Matrix (ISO 17034) |  | 20 weeks from contract effectiveness | 10 No. |  |  |  |  |
| 142 | pH Buffer Standards (pH 4.00 / 7.00 / 10.00, NIST-traceable) |  | 20 weeks from contract effectiveness | 24 No. |  |  |  |  |
| 143 | Karl Fischer Reagent (Composite grade, 1 L) |  | 20 weeks from contract effectiveness | 10 No. |  |  |  |  |
| 144 | Argon Gas – Purity ≥99.995% (8500 L cylinder) |  | 20 weeks from contract effectiveness | 4 No. |  |  |  |  |
| 145 | Nitrogen Gas – Purity ≥99.999% (8500 L cylinder) |  | 20 weeks from contract effectiveness | 4 No. |  |  |  |  |
| 146 | Acetylene Gas – Dissolved ≥99.5% (8 kg cylinder) |  | 20 weeks from contract effectiveness | 4 No. |  |  |  |  |
| 147 | Borosilicate Glassware Sets (Beakers, Flasks, Cylinders, Volumetric) |  | 20 weeks from contract effectiveness | 10 No. |  |  |  |  |
| 148 | Centrifuge Tubes (50 mL + 15 mL PP), DNase/RNase-free |  | 20 weeks from contract effectiveness | 20 No. |  |  |  |  |
| 149 | Microcentrifuge Tubes (1.5 mL + 2.0 mL), PCR-clean |  | 20 weeks from contract effectiveness | 30 No. |  |  |  |  |
| 150 | PCR Plates (96-well, 0.2 mL, thin-wall) |  | 20 weeks from contract effectiveness | 30 No. |  |  |  |  |
| 151 | Pipette Tips (10 µL, 200 µL, 1000 µL) – Filtered & Unfiltered |  | 20 weeks from contract effectiveness | 60 No. |  |  |  |  |
| 152 | Filter Papers – Grades 1, 4, 41, 42 (100/box) |  | 20 weeks from contract effectiveness | 20 No. |  |  |  |  |
| 153 | Membrane Filters – 0.2 µm & 0.45 µm (47 mm, 50/box) |  | 20 weeks from contract effectiveness | 20 No. |  |  |  |  |
| 154 | Disposable Nitrile Gloves – Powder-free (100/box, S/M/L) |  | 20 weeks from contract effectiveness | 200 No. |  |  |  |  |
| 155 | Lab Coats – White, Knee-length |  | 20 weeks from contract effectiveness | 60 No. |  |  |  |  |
| 156 | FFP2 / N95 Masks (Non-valve, EN 149/NIOSH) |  | 20 weeks from contract effectiveness | 200 No. |  |  |  |  |
| 157 | Fat Extraction Thimbles – Cellulose (33×80 mm, 100/box) |  | 20 weeks from contract effectiveness | 20 No. |  |  |  |  |
| 158 | Rapid Test Kits – Livestock Disease Panel (Annual, 6 types) |  | 20 weeks from contract effectiveness | 80 No. |  |  |  |  |
| 159 | Rapid Test Kits – Meat & Seafood Safety (Annual, 8 types) |  | 20 weeks from contract effectiveness | 100 No. |  |  |  |  |
|  |  |  |  |  |  |  | Total Price |  |
| Name of Bidder [insert complete name of Bidder] Signature of Bidder [signature of person signing the Bid] Date [Insert Date] |  |  |  |  |  |  |  |  |

| Price Schedule: Goods Manufactured Outside the Purchaser’s Country, already imported* |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  | (Group C Bids, Goods already imported) Currencies in accordance with ITB 15 |  |  |  |  |  | Date:_________________________ RFB No: _____________________ Alternative No: ________________ Page N ______ of ______ |  |  |
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 |
| Line Item N | Description of Goods | Country of Origin | Delivery Date as defined by Incoterms | Quantity and physical unit | Unit price including Custom Duties and Import Taxes paid, in accordance with ITB 14.8(c)(i) | Custom Duties and Import Taxes paid per unit in accordance with ITB 14.8(c)(ii) , [to be supported by documents] | Unit Price net of custom duties and import taxes, in accordance with ITB 14.8 (c) (iii) (Col. 6 minus Col.7) | Price per line item net of Custom Duties and Import Taxes paid, in accordance with ITB 14.8(c)(i) (Col. 58) | Price per line item for inland transportation and other services required in the Purchaser’s Country to convey the goods to their final destination, as specified in BDS in accordance with ITB 14.8 (c)(v) | Sales and other taxes paid or payable per item if Contract is awarded (in accordance with ITB 14.8(c)(iv) | Total Price per line item (Col. 9+10) |
| [insert number of the item] | [insert name of Goods] | [insert country of origin of the Good] | [insert quoted Delivery Date] | [insert number of units to be supplied and name of the physical unit] | [insert unit price per unit] | [insert custom duties and taxes paid per unit] | [insert unit price net of custom duties and import taxes] | [ insert price per line item net of custom duties and import taxes] | [insert price per line item for inland transportation and other services required in the Purchaser’s Country] | [insert sales and other taxes payable per item if Contract is awarded] | [insert total price per line item] |
|  |  |  |  |  |  |  |  |  |  | Total Bid Price |  |
| Name of Bidder [insert complete name of Bidder] Signature of Bidder [signature of person signing the Bid] Date [insert date] |  |  |  |  |  |  |  |  |  |  |  |

* [For previously imported Goods, the quoted price shall be distinguishable from the original import value of these Goods declared to customs and shall include any rebate or mark-up of the local agent or representative and all local costs except import duties and taxes, which have been and/or have to be paid by the Purchaser. For clarity the Bidders are asked to quote the price including import duties, and additionally to provide the import duties and the price net of import duties which is the difference of those values.]

Price Schedule: Goods Manufactured in the Purchaser’s Country

| Purchaser’s Country ______________________ |  |  |  | (Group A and B Bids) Currencies in accordance with ITB 15 |  |  |  | Date:_________________________ RFB No: _____________________ Alternative No: ________________ Page N ______ of ______ |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 |
| Line Item N | Description of Goods | Delivery Date as defined by Incoterms | Quantity and physical unit | Unit price EXW | Total EXW price per line item (Col. 45) | Price per line item for inland transportation and other services required in the Purchaser’s Country to convey the Goods to their final destination | Cost of local labor, raw materials and components from with origin in the Purchaser’s Country % of Col. 5 | Sales and other taxes payable per line item if Contract is awarded (in accordance with ITB 14.8(a)(ii) | Total Price per line item (Col. 6+7) |
| [insert number of the item] | [insert name of Good] | [insert quoted Delivery Date] | [insert number of units to be supplied and name of the physical unit] | [insert EXW unit price] | [insert total EXW price per line item] | [insert the corresponding price per line item] | [Insert cost of local labor, raw material and components from within the Purchase’s country as a % of the EXW price per line item] | [insert sales and other taxes payable per line item if Contract is awarded] | [insert total price per item] |
|  |  |  |  |  |  |  |  | Total Price |  |
| Name of Bidder [insert complete name of Bidder] Signature of Bidder [signature of person signing the Bid] Date [insert date] |  |  |  |  |  |  |  |  |  |

Price and Completion Schedule: Related Services

Currencies in accordance with ITB 15

| 1 | 2 | 3 | 4 | 5 | 6 | 7 |
| --- | --- | --- | --- | --- | --- | --- |
| Service Nº | Description of Services | Country of Origin | Completion Date | Quantity and physical unit | Unit price | Total Price per Service (Col. 5×6) |
| 1 | Installation, commissioning and qualification (IQ/OQ/PQ) of all instruments and equipment |  | Within eight (8) weeks of delivery of the respective equipment | 1 Lot |  |  |
| 2 | Instrument Training – HPLC (3-day, on-site) |  | Within eight (8) weeks of completion of installation and commissioning | 1 course |  |  |
| 3 | Instrument Training – GC-MS (3-day, on-site) |  | Within eight (8) weeks of completion of installation and commissioning | 1 course |  |  |
| 4 | Instrument Training – AAS (2-day, on-site) |  | Within eight (8) weeks of completion of installation and commissioning | 1 course |  |  |
| 5 | Instrument Training – qPCR (2-day, on-site) |  | Within eight (8) weeks of completion of installation and commissioning | 1 course |  |  |
| 6 | LIMS Administrator + User Training (5-day) |  | Within eight (8) weeks of completion of installation and commissioning | 1 course |  |  |
| 7 | Proficiency Testing – Annual Enrolment (ISO/IEC 17043-accredited provider) |  | First year of operation; enrolment within twelve (12) months of contract effectiveness | 10 rounds |  |  |
| 8 | External Calibration Services (ISO/IEC 17025-accredited laboratory) |  | First year of operation, within twelve (12) months of completion of commissioning | 15 events |  |  |
| Total Related Services price |  |  |  |  |  |  |

Name of Bidder [insert complete name of Bidder] Signature of Bidder [signature of person signing the Bid] Date [Insert Date]

| Price and Completion Schedule - Related Services |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Currencies in accordance with ITB 15 |  |  |  | Date:_________________________ RFB No: _____________________ Alternative No: ________________ Page N ______ of ______ |  |
| 1 | 2 | 3 | 4 | 5 | 6 | 7 |
| Service N | Description of Services (excludes inland transportation and other services required in the Purchaser’s Country to convey the goods to their final destination) | Country of Origin | Delivery Date at place of Final destination | Quantity and physical unit | Unit price | Total Price per Service (Col. 5*6 or estimate) |
| [insert number of the Service ] | [insert name of Services] | [insert country of origin of the Services] | [insert delivery date at place of final destination per Service] | [insert number of units to be supplied and name of the physical unit] | [insert unit price per item] | [insert total price per item] |
|  |  |  |  | Total Bid Price |  |  |
| Name of Bidder [insert complete name of Bidder] Signature of Bidder [signature of person signing the Bid] Date [insert date] |  |  |  |  |  |  |

Section V - Eligible Countries

Eligibility for the Provision of Goods, Works and Non-Consulting Services in 
Bank-Financed Procurement

In reference to ITB 4.8 and ITB 5.1, for the information of the Bidders, at the present time firms, goods and services from the following countries are excluded from this Bidding process:

Under ITB 4.8 (a) and ITB 5.1: “none”

Under ITB 4.8(b) and ITB 5.1: “none”

Section VI - Fraud and Corruption

(Section VI shall not be modified)

- Purpose

  - The Bank’s Anti-Corruption Guidelines and this annex apply with respect to procurement under Bank Investment Project Financing operations.

- Requirements

- The Bank requires that Borrowers (including beneficiaries of Bank financing); bidders (applicants/proposers), consultants, contractors and suppliers; any sub-contractors, sub-consultants, service providers or suppliers; any agents (whether declared or not); and any of their personnel, observe the highest standard of ethics during the procurement process, selection and contract execution of Bank-financed contracts, and refrain from Fraud and Corruption.

- To this end, the Bank:

- Defines, for the purposes of this provision, the terms set forth below as follows:

- “corrupt practice” is the offering, giving, receiving, or soliciting, directly or indirectly, of anything of value to influence improperly the actions of another party;

- “fraudulent practice” is any act or omission, including misrepresentation, that knowingly or recklessly misleads, or attempts to mislead, a party to obtain financial or other benefit or to avoid an obligation;

- “collusive practice” is an arrangement between two or more parties designed to achieve an improper purpose, including to influence improperly the actions of another party;

- “coercive practice” is impairing or harming, or threatening to impair or harm, directly or indirectly, any party or the property of the party to influence improperly the actions of a party;

- “obstructive practice” is:

- deliberately destroying, falsifying, altering, or concealing of evidence material to the investigation or making false statements to investigators in order to materially impede a Bank investigation into allegations of a corrupt, fraudulent, coercive, or collusive practice; and/or threatening, harassing, or intimidating any party to prevent it from disclosing its knowledge of matters relevant to the investigation or from pursuing the investigation; or

- acts intended to materially impede the exercise of the Bank’s inspection and audit rights provided for under paragraph 2.2 e. below.

- Rejects a proposal for award if the Bank determines that the firm or individual recommended for award, any of its personnel, or its agents, or its sub-consultants, sub-contractors, service providers, suppliers and/ or their employees, has, directly or indirectly, engaged in corrupt, fraudulent, collusive, coercive, or obstructive practices in competing for the contract in question;

- In addition to the legal remedies set out in the relevant Legal Agreement, may take other appropriate actions, including declaring misprocurement, if the Bank determines at any time that representatives of the Borrower or of a recipient of any part of the proceeds of the loan engaged in corrupt, fraudulent, collusive, coercive, or obstructive practices during the procurement process, selection and/or execution of the contract in question,  without the Borrower having taken timely and appropriate action satisfactory to the Bank to address such practices when they occur, including by failing to inform the Bank in a timely manner at the time  they knew of the practices;

- Pursuant to the Bank’s Anti-Corruption Guidelines, and in accordance with the Bank’s prevailing sanctions policies and procedures, may sanction a firm or individual, either indefinitely or for a stated period of time, including by publicly declaring such firm or individual ineligible (i) to be awarded or otherwise benefit from a Bank-financed contract, financially or in any other manner; (ii) to be a nominated sub-contractor, consultant, manufacturer or supplier, or service provider of an otherwise eligible firm being awarded a Bank-financed contract; and (iii) to receive the proceeds of any loan made by the Bank or otherwise to participate further in the preparation or implementation of any Bank-financed project;

- Requires that a clause be included in bidding/request for proposals documents and in contracts financed by a Bank loan, requiring (i) bidders (applicants/proposers), consultants, contractors, and suppliers, and their sub-contractors, sub-consultants, service providers, suppliers, agents, personnel, permit the Bank to inspect all accounts, records and other documents relating to the procurement process, selection and/or contract execution, and to have them audited by auditors appointed by the Bank.

PART 2 – Supply Requirements

Section VII - Schedule of Requirements

Contents

1.	List of Goods and Delivery Schedule	95

2.	List of Related Services and Completion Schedule	96

3. Technical Specifications	97

4. Drawings	100

5. Inspections and Tests	101

| - List of Goods and Delivery Schedule “Bidder’s offered Delivery date” to be filled by the Bidder |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| A. Analytical Chemistry Equipment |  |  |  |  |  |  |  |
| Item No. | Description of Goods | Quantity | Unit | Final Destination | Delivery Period Start Date | Required Delivery Period | Bidder’s offered Delivery date |
| 1 | LC-MS/MS System (Triple Quadrupole) | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 2 | GC-MS/MS System (Triple Quadrupole) | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 3 | HPLC System (Quaternary Gradient, DAD + FLD) | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 4 | GC-MS System (Single Quadrupole + FID/ECD) | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 5 | AAS – Flame + Graphite Furnace (Zeeman) | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 6 | FTIR Spectrometer (Benchtop ATR) | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 7 | UV/Vis Spectrophotometer (Double-Beam, Scanning) | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 8 | NIR Analyzer (Rapid Compositional, Benchtop) | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 9 | Ion Chromatograph (Suppressed Conductivity) | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 10 | Karl Fischer Titrator (Coulometric) | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 11 | HPLC Column – C18 Reverse Phase (150×4.6 mm, 5 µm) | 6 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 12 | HPLC Column – C8 (150×4.6 mm, 5 µm) | 4 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 13 | HPLC Column – HILIC (100×2.1 mm, 1.8 µm) | 4 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 14 | HPLC Column – Phenyl-Hexyl (150×4.6 mm, 5 µm) | 4 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 15 | HPLC Guard Columns (C18, 12.5×4.6 mm) – Pack of 3 | 12 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 16 | GC Capillary Column – DB-5 equivalent (30 m × 0.25 mm × 0.25 µm) | 4 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 17 | GC Capillary Column – Mid-polarity (30 m × 0.25 mm × 0.25 µm) | 2 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 18 | GC Capillary Column – Wax-phase (30 m × 0.25 mm × 0.25 µm) | 2 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 19 | SPE Cartridges – C18 (200 mg / 3 mL), Box 50 | 20 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 20 | SPE Cartridges – Mixed-mode Dispersive (QuEChERS), Box 50 | 20 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 21 | Immunoaffinity Columns – Aflatoxin (25/box) | 8 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 22 | Immunoaffinity Columns – Ochratoxin A (25/box) | 6 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 23 | Immunoaffinity Columns – Fumonisin B1/B2 (25/box) | 6 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| B. Microbiology Equipment |  |  |  |  |  |  |  |
| 24 | Biosafety Cabinet – Class II A2 (1.2 m) | 4 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 25 | Autoclave / Steam Sterilizer (100 L, Pre-Vacuum) | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 26 | Anaerobic Incubator / Workstation | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 27 | Incubator – Temperature Controlled (100–400 L) | 4 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 28 | Real-Time PCR System (96-well, 4–6 channel) | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 29 | Deep Freezer –80°C (Ultra-Low, 400 L) | 2 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 30 | Refrigerator (2–8°C, 250 L, Lab-Grade) | 4 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 31 | Automated Microplate Reader (ELISA) | 3 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 32 | Colony Counter (Digital, Image Analysis) | 5 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 33 | Fluorescence Microscope (Upright) | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 34 | Impedance / ATP Microbial Analyzer | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 35 | Culture Media – Annual Supply (Selective + Non-selective, 14 types) | 150 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 36 | Reference Cultures – Freeze-dried (8 species, certified) | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| C. Sample Preparation and General Laboratory Equipment |  |  |  |  |  |  |  |
| 37 | Analytical Balance (High-Precision, 220 g / 0.01 mg) | 2 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 38 | Precision Balance (50 kg / 1 g) | 2 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 39 | pH Meter (Bench-top, Multi-parameter, IP54) | 4 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 40 | Moisture Analyzer (Halogen/IR) | 3 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 41 | Refractometer (Digital, Peltier-controlled) | 2 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 42 | Homogenizer / Stomacher – 400 mL (Paddle Blender) | 2 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 43 | Stomacher – High-Volume (3500 mL) | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 44 | High-Speed Centrifuge – Refrigerated (0–25,000 rpm) | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 45 | Microcentrifuge (Benchtop, 0–16,000 rpm) | 2 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 46 | Water Bath / Heating Bath (5–99°C) | 3 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 47 | Dry Bath / Block Heater (Ambient+5 to 130°C) | 3 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 48 | Orbital Shaker / Incubator Shaker | 2 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 49 | Vortex Mixer (Variable 100–3200 rpm) | 4 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 50 | Magnetic Stirrer / Hot Plate (Combined) | 6 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 51 | Vacuum Oven (100 L, Ambient–250°C) | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 52 | Muffle Furnace – Ash Determination (1200°C max) | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 53 | Drying Oven – Forced Air (Ambient to 300°C) | 3 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 54 | Ultrasonic Cleaner (20 L, 40 kHz) | 2 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 55 | Water Purification System (Type I + Type II) | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 56 | Turbidimeter / Nephelometer (0–4000 NTU) | 3 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 57 | Dissolved Oxygen / Multi-parameter Meter | 2 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 58 | Water Activity Meter (0–1.00 aw) | 2 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 59 | Texture Analyzer (0.01–250 N) | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 60 | Colorimeter (Meat/Food Colour) | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 61 | Pipette Set – Single Channel (6 volume ranges) | 6 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 62 | Multi-Channel Pipette (8 & 12-channel, 50–300 µL) | 4 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| D. Infrastructure and Support Equipment |  |  |  |  |  |  |  |
| 63 | UPS System – Online, 10 kVA (for LC-MS/MS, HPLC, AAS) | 4 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 64 | UPS System – Online, 3 kVA (for PCR, Freezers, Incubators) | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 65 | Power Line Conditioner / Voltage Stabilizer (5 kVA) | 6 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 66 | HVAC – Dedicated Lab Precision Air Handling Unit | 2 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 67 | Fume Hood – Ducted Chemical (1.2–1.5 m) | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 68 | Sample Reception Area Setup | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 69 | Cold Storage Tracking System (Integrated with LIMS) | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 70 | Sample Storage Racks – Traceable (Cryogenic + Ambient) | 20 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 71 | Refrigerator -20°C (Intermediate Sample Storage, 200 L) | 2 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 72 | Cold Chain Maintenance (Portable Coolers, PCM Ice Packs) | 10 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 73 | Data Logger – Temperature/Humidity (USB, Alarm) | 10 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 74 | Emergency Shower & Eyewash Station – Plumbed (ANSI Z358.1) | 3 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 75 | Portable Eyewash Station (Self-contained, gravity-fed) | 4 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 76 | Chemical Spill Kit (Acid + Base + Solvent) | 8 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 77 | Biohazard Waste Containers (Autoclavable Bags + Sharps) | 20 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 78 | Chemical Waste Storage Cabinet (Acid-Resistant, 90 L) | 4 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 79 | CO₂ Fire Extinguisher (5 kg, EN 3) | 10 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 80 | Photoelectric Smoke Detector (IP44, Lab-grade) | 8 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 81 | First Aid Kit (Laboratory Grade, 50-person) | 4 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| E. Enabling Systems (IT, LIMS and Calibration Support Equipment) |  |  |  |  |  |  |  |
| 82 | Barcode Label Printer (Lab-grade, 300 dpi, Thermal) | 4 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 83 | Barcode Scanner (USB + Wireless, 1D/2D) | 8 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 84 | Role-Based Access Control System (LIMS + Instrument Access) | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 85 | Environmental Monitoring Sensors (Wireless, Temp/RH) | 20 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 86 | Certified Weight Set – OIML Class E2/F1 (1 mg to 2 kg) | 3 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 87 | Pipette Calibration System – Gravimetric (ISO 8655) | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 88 | Temperature Calibration – Dry Block Calibrator + Reference Thermometer | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 89 | Calibration Thermometer – Handheld, 4-channel, Pt100 | 4 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 90 | Laboratory Server / NAS Backup System | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 91 | Network Infrastructure (Managed Switch + Lab Wi-Fi) | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 92 | Analyst Workstation – Instrument-Dedicated (Desktop) | 4 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 93 | Lab Analyst PC – LIMS & Reporting (Desktop) | 6 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 94 | Lab Manager / QA Officer Laptop | 2 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 95 | Network Laser Printer – A4/A3 (Mono + Colour, Duplex) | 2 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 96 | UPS – Line-Interactive (1 kVA) for Workstations & PCs | 12 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 97 | External Hard Drive – Encrypted Backup (4 TB, USB 3.0) | 6 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| F. Livestock and Feed Testing Equipment |  |  |  |  |  |  |  |
| 98 | Milk Analyzer – Multi-parameter (Fat/Protein/Lactose/SNF/Conductivity) | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 99 | Hematology Analyzer – Automated CBC (Livestock Blood Panels) | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 100 | Veterinary Biochemistry Analyzer (Photometric, Automated) | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 101 | Portable Milk Tester – Field Grade | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 102 | Gerber Centrifuge (12-position, ISO 2446) | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 103 | Portable Livestock Disease Rapid Test Reader | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 104 | Stereomicroscope (Feed Adulteration Examination) | 2 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 105 | Feed Particle Sieve Analyser (Vibrating Sieve Tower) | 1 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 106 | ELISA Kit – Brucellosis (B. abortus + B. melitensis) | 5 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 107 | ELISA Kit – Foot & Mouth Disease (FMD, Serotypes O/A/SAT) | 5 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 108 | ELISA Kit – African Swine Fever (ASFV Antigen + Antibody) | 10 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 109 | ELISA Kit – Peste des Petits Ruminants (PPR, cELISA) | 10 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 110 | ELISA Kit – Rift Valley Fever (RVF, IgM/IgG) | 5 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 111 | ELISA Kit – Bovine Tuberculosis (IFN-gamma, 24-well) | 10 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 112 | ELISA Kit – Avian Influenza H5/H7 (Competitive ELISA) | 10 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 113 | Lateral Flow Strip – Beta-Lactam Antibiotics (Milk + Tissue) | 4 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 114 | Lateral Flow Strip – Tetracyclines (4-in-1 card) | 4 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 115 | Lateral Flow Strip – Chloramphenicol (Banned Substance) | 2 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 116 | Lateral Flow Strip – Streptomycin / Aminoglycosides | 2 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 117 | Lateral Flow Strip – Clenbuterol (Beta-agonist Growth Promoter) | 4 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 118 | Lateral Flow Strip – Aflatoxin M1 in Milk | 4 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 119 | Blood Collection Tubes – EDTA (10 mL, 1000/box) | 5 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 120 | Blood Collection Tubes – Plain/SST (10 mL, 1000/box) | 5 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 121 | Vacutainer Needles + Holders (21G × 1.5", 100/box) | 30 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 122 | Animal Ear Tag Labels + Barcode Stickers (1000/roll) | 10 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 123 | Feed Sampling Probe (Stainless Steel, 1.5 m) | 5 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 124 | Feed CRM – Proximate Analysis (Protein/Fat/Moisture/Ash/Fibre) | 10 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 125 | Milk CRM – Compositional Analysis (Fat/Protein/Lactose) | 6 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| G. Consumables and Reagents (First-Year Supply) |  |  |  |  |  |  |  |
| 126 | Acetonitrile – HPLC Grade (2.5 L) | 20 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 127 | Methanol – HPLC Grade (2.5 L) | 20 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 128 | Hexane – HPLC Grade (2.5 L) | 15 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 129 | Acetone – HPLC Grade (2.5 L) | 15 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 130 | Dichloromethane – HPLC Grade (2.5 L) | 10 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 131 | Ethanol – AR Grade (2.5 L) | 15 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 132 | Acetic Acid – Glacial AR (2.5 L) | 10 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 133 | Nitric Acid – Trace Metal Grade (500 mL) | 8 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 134 | Hydrochloric Acid – AR Grade (2.5 L) | 10 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 135 | Sulfuric Acid – AR Grade (2.5 L) | 10 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 136 | Sodium Hydroxide – AR Pellets (1 kg) | 10 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 137 | Heavy Metal Standards – 1000 mg/L (8 elements) | 24 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 138 | Pesticide Reference Standard Mix (100+ compounds) | 6 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 139 | Veterinary Drug Standard Mix (50–100 compounds) | 6 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 140 | Mycotoxin Standards (AFB1/B2/G1/G2, OTA, Fumonisin B1/B2) | 12 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 141 | Certified Reference Material – Meat/Fish Matrix (ISO 17034) | 10 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 142 | pH Buffer Standards (pH 4.00 / 7.00 / 10.00, NIST-traceable) | 24 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 143 | Karl Fischer Reagent (Composite grade, 1 L) | 10 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 144 | Argon Gas – Purity ≥99.995% (8500 L cylinder) | 4 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 145 | Nitrogen Gas – Purity ≥99.999% (8500 L cylinder) | 4 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 146 | Acetylene Gas – Dissolved ≥99.5% (8 kg cylinder) | 4 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 147 | Borosilicate Glassware Sets (Beakers, Flasks, Cylinders, Volumetric) | 10 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 148 | Centrifuge Tubes (50 mL + 15 mL PP), DNase/RNase-free | 20 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 149 | Microcentrifuge Tubes (1.5 mL + 2.0 mL), PCR-clean | 30 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 150 | PCR Plates (96-well, 0.2 mL, thin-wall) | 30 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 151 | Pipette Tips (10 µL, 200 µL, 1000 µL) – Filtered & Unfiltered | 60 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 152 | Filter Papers – Grades 1, 4, 41, 42 (100/box) | 20 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 153 | Membrane Filters – 0.2 µm & 0.45 µm (47 mm, 50/box) | 20 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 154 | Disposable Nitrile Gloves – Powder-free (100/box, S/M/L) | 200 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 155 | Lab Coats – White, Knee-length | 60 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 156 | FFP2 / N95 Masks (Non-valve, EN 149/NIOSH) | 200 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 157 | Fat Extraction Thimbles – Cellulose (33×80 mm, 100/box) | 20 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 158 | Rapid Test Kits – Livestock Disease Panel (Annual, 6 types) | 80 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |
| 159 | Rapid Test Kits – Meat & Seafood Safety (Annual, 8 types) | 100 | No. | SoBS – NRL, Mogadishu | Contract effectiveness | 20 weeks from contract effectiveness |  |

| - List of Related Services and Completion Schedule |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Service | Description of Service | Quantity1 | Physical Unit | Place where Services shall be performed | Final Completion Date(s) of Services |
| 1 | Installation, commissioning and qualification (IQ/OQ/PQ) of all instruments and equipment listed in the List of Goods, including site preparation verification, utility connections, calibration certificates and acceptance testing | 1 | Lot | Somali Bureau of Standards (SoBS) – National Reference Laboratory (NRL), Mogadishu | Within eight (8) weeks of delivery of the respective equipment |
| 2 | Instrument Training – HPLC (3-day, on-site). 3-day on-site; 4–6 analysts; data system software; column selection, gradient optimization, system suitability testing | 1 | course | Somali Bureau of Standards (SoBS) – National Reference Laboratory (NRL), Mogadishu | Within eight (8) weeks of completion of installation and commissioning of the respective instrument(s) |
| 3 | Instrument Training – GC-MS (3-day, on-site). 3-day on-site; EU EN 15662 QuEChERS methods; data analysis software; GC maintenance procedures | 1 | course | Somali Bureau of Standards (SoBS) – National Reference Laboratory (NRL), Mogadishu | Within eight (8) weeks of completion of installation and commissioning of the respective instrument(s) |
| 4 | Instrument Training – AAS (2-day, on-site). 2-day on-site; Codex/EU MRL methods; graphite furnace optimization; lamp alignment; hydride generation techniques | 1 | course | Somali Bureau of Standards (SoBS) – National Reference Laboratory (NRL), Mogadishu | Within eight (8) weeks of completion of installation and commissioning of the respective instrument(s) |
| 5 | Instrument Training – qPCR (2-day, on-site). 2-day on-site; primer/probe design; ISO 20837 food pathogen methods; instrument software; 4–6 analysts | 1 | course | Somali Bureau of Standards (SoBS) – National Reference Laboratory (NRL), Mogadishu | Within eight (8) weeks of completion of installation and commissioning of the respective instrument(s) |
| 6 | LIMS Administrator + User Training. 5-day training; 2 system admins and 10 end-users; sample management, QC charting, report generation, audit trail review | 1 | course | Somali Bureau of Standards (SoBS) – National Reference Laboratory (NRL), Mogadishu | Within eight (8) weeks of completion of installation and commissioning of the respective instrument(s) |
| 7 | Proficiency Testing – Annual Enrolment (10 rounds). Internationally accredited PT scheme provider; schemes: mycotoxins, pesticide residues, heavy metals, nutritional composition, microbiology; 10 rounds/year | 10 | round | SoBS – NRL, Mogadishu (samples via accredited PT scheme provider) | First year of laboratory operation; enrolment completed within twelve (12) months of contract effectiveness |
| 8 | External Calibration Services. ISO 17025 accredited calibration laboratory; services: balances, instruments (HPLC/GC IQ/OQ/PQ), temperature mapping, pipette verification; 15 calibration events | 15 | event | Somali Bureau of Standards (SoBS) – National Reference Laboratory (NRL), Mogadishu | First year of laboratory operation, within twelve (12) months of completion of commissioning |
| 9 | Pre-delivery site-readiness assessment and interface confirmation, including access, floor/bench loading, electrical, earthing, HVAC, exhaust/fume hoods, gases, water/drainage, ICT/networking, cold-chain and safety interfaces | 1 | service event | Somali Bureau of Standards (SoBS) – National Reference Laboratory (NRL), Mogadishu | Within twenty-one (21) days after Contract signature and, in any event, before shipment of the Goods |
| 10 | Handover documentation package, including manuals, SOP templates, certificates, warranty documents, software licenses and administrator credentials, data ownership/export documentation, backup/restore evidence and as-installed records, and environmental, social, health and safety (ESHS) close-out records (waste and packaging removal evidence, Safety Data Sheet (SDS) register handover and incident log) | 1 | documentation package | Somali Bureau of Standards (SoBS) – National Reference Laboratory (NRL), Mogadishu | Before Operational Acceptance |
| 11 | Cybersecurity hardening and LIMS/interface handover, including access-control setup, patching baseline, audit/logging configuration, backup/restore test and breach-notification procedure | 1 | service package | Somali Bureau of Standards (SoBS) – National Reference Laboratory (NRL), Mogadishu | Before Operational Acceptance |
| 12 | Laboratory Safety and Hazardous-Materials Handling Training (2-day, on-site). 2-day on-site; up to 10 SoBS participants; chemical hazards and Safety Data Sheet (SDS) use, gas-cylinder safety, spill response, biosafety basics, waste segregation, PPE, emergency shower/eyewash use, fire safety | 1 | course | Somali Bureau of Standards (SoBS) – National Reference Laboratory (NRL), Mogadishu | Before Operational Acceptance |

Note: Item-specific calibration certificates and instrument-specific training already required under Goods line items shall not be double-priced under Related Services.

3. Technical Specifications

“Summary of Technical Specifications. The Goods and Related Services shall comply with following Technical Specifications and Standards:

| Item No | Name of Goods or Related Service | Technical Specifications and Standards |
| --- | --- | --- |
| 1 | Analytical Chemistry Equipment – Items 1–23 | See detailed specifications below |
| 2 | Microbiology Equipment – Items 24–36 | See detailed specifications below |
| 3 | Sample Preparation and General Laboratory Equipment – Items 37–62 | See detailed specifications below |
| 4 | Infrastructure and Support Equipment – Items 63–81 | See detailed specifications below |
| 5 | Enabling Systems (IT, LIMS and Calibration Support Equipment) – Items 82–97 | See detailed specifications below |
| 6 | Livestock and Feed Testing Equipment – Items 98–125 | See detailed specifications below |
| 7 | Consumables and Reagents (First-Year Supply) – Items 126–159 | See detailed specifications below |
| 8 | Related Services – installation, commissioning and IQ/OQ/PQ qualification; instrument and LIMS training; proficiency testing enrolment; external calibration services; pre-delivery site-readiness assessment; handover documentation package; cybersecurity hardening and LIMS/interface handover; laboratory safety and hazardous-materials handling training | See List of Related Services and Completion Schedule, and service requirements below |

Detailed Technical Specifications and Standards

Schedule of Requirement

Supply, Delivery, Installation and Commissioning of Laboratory Equipment for the National Reference Laboratory (NRL) of the Somali Bureau of Standards (SoBS)

Detailed Technical Specifications for Laboratory Equipment, Consumables and Related Services

The detailed technical specifications below state the Purchaser’s minimum requirements. Where a reference standard or classification is mentioned, it is for the purpose of establishing the required performance and quality level; equivalent or superior products fully meeting the stated specifications are acceptable.

Bidders shall complete the “Bidder’s Response” column for every line item, stating “Comply” (or describing any deviation) and providing a precise cross-reference to the manufacturer’s published technical literature (datasheet or catalogue page) submitted with the Bid. Failure to provide a substantiated response for any item may render the Bid nonresponsive.

All equipment shall be new, unused and of current manufacture; supplied complete with standard accessories, installation kits, start-up consumables and first-fill reagents required for commissioning; suitable for operation at the site’s electrical supply (220–240 V, 50 Hz) with all necessary plugs, adapters, stabilizers or transformers included; and shall carry relevant international certification (CE or equivalent) and comply with the ISO/IEC/EN standards cited in the specifications, where applicable.

Site conditions and interface requirements.

The Goods shall be installed at the Somali Bureau of Standards (SoBS) – National Reference Laboratory (NRL), Mogadishu. Site conditions shall be confirmed through the pre-delivery site-readiness assessment (Related Service No. 9 of the List of Related Services).

The Supplier shall review, verify and plan against the following interfaces:

| Interface area | Supplier responsibility | Evidence required in Bid / before mobilization |
| --- | --- | --- |
| Access route, unloading and internal movement | Review and verify site access; provide lifting/offloading method, equipment protection and logistics plan. | Delivery/logistics plan and risk assessment. |
| Electrical supply, earthing and power backup (site supply 220–240 V, 50 Hz) | Identify loads, plug types, earthing and UPS/power-conditioning needs for each item. | Electrical interface matrix and manufacturer requirements. |
| HVAC, temperature, humidity and exhaust | State heat loads, ventilation/exhaust/fume-hood requirements and operating tolerances. | HVAC/exhaust interface schedule. |
| Laboratory gases and cylinders | Specify gas purity, cylinder/regulator/manifold requirements, storage and safety controls. | Gas interface and safety plan. |
| Water, purified water and drainage | Identify water quality, flow, drainage and waste-water interface requirements. | Water/drainage interface matrix. |
| IT / network / LIMS / cybersecurity | Identify network, device, LIMS/interface, user-access, backup and cybersecurity requirements. | LIMS/ICT/cybersecurity plan. |
| Cold-chain and hazardous materials storage | Manage storage temperature, SDS, cold-chain monitoring, segregation and emergency controls. | SDS register, cold-chain and hazardous-material plan. |
| Waste, packaging and e-waste | Remove and lawfully manage Supplier-generated packaging, installation waste, replaced parts, batteries and e-waste. | Waste/e-waste management and close-out records. |

Hazardous and temperature-controlled items

For chemicals, reagents, reference standards, gases, refrigerants, batteries or other hazardous or temperature-controlled items supplied under the Contract, the Supplier shall provide Safety Data Sheets, certificates of analysis (where applicable), expiry dates, and instructions for safe handling, transport, storage, spill response and disposal, together with cold-chain temperature monitoring records for temperature-controlled items. The Supplier shall not dispose of any hazardous waste, chemical waste or electronic waste within Somalia. All hazardous waste, expired or damaged chemicals, packaging, crating and replaced components shall be removed from the Purchaser’s premises, segregated, packaged, labelled and transported in accordance with applicable regulations, and shall be taken back by the Supplier or transferred to a facility lawfully permitted to manage such waste; the Supplier shall provide the Purchaser with evidence of lawful management for each consignment removed.

Cybersecurity

For LIMS/QMS software, networked instruments, computers, servers/NAS, data storage and remote-support components, the Supplier shall implement reasonable cybersecurity controls, including user access control, role-based permissions, audit trails, malware protection, secure configuration, patching and update arrangements, and backup/restore testing, and shall hand over all administrator credentials and licence documentation to the Purchaser. Cybersecurity does not apply to purely standalone, non-networked equipment except to the extent software, firmware or removable media is used.

Environmental, Social, Health and Safety (ESHS). ESHS requirements form part of these Technical Specifications. The Supplier shall comply with the World Bank Environmental and Social Standards applicable to the Project and with the DRIVE Project E&S instruments — the Environmental and Social Commitment Plan (ESCP), the Environmental and Social Management Framework (ESMF), the Stakeholder Engagement Plan (SEP) and the E&S procedures of the Project Implementation Manual — as updated or supplemented by site-specific requirements issued by the Purchaser. Before mobilization to site, the Supplier shall submit and implement an ESHS/OHS management plan acceptable to the Purchaser covering transport, offloading, lifting and installation method statements, PPE, site access, fire safety, spill prevention, emergency response, a worker Code of Conduct (including SEA/SH prevention), and incident reporting. The Supplier shall notify the Purchaser immediately, and in any event within twenty-four (24) hours, of any serious incident related to the Contract (including injury, fire, spill, exposure, security incident, SEA/SH allegation, environmental release, cold-chain failure or cybersecurity incident), followed by a written incident report within forty-eight (48) hours. The Supplier shall remove or ensure lawful and environmentally sound management of packaging waste, installation waste, replaced parts, batteries and e-waste generated by its activities.

Worker Code of Conduct

The Supplier shall ensure that all its personnel, and its subcontractors’ personnel, engaged in the delivery, offloading, installation or commissioning of the Goods at the Project Site read and sign the following Code of Conduct before commencing work on site, and shall keep the signed copies available for the Purchaser’s inspection:

(1) comply with all site rules, safety instructions and PPE requirements;

(2) treat all persons with respect; no discrimination, harassment, violence or abusive behaviour;

(3) no sexual exploitation or abuse and no sexual harassment, including no sexual activity with any person under the age of 18, no exchange of money, employment, goods or services for sex, and no unwelcome sexual advances; any SEA/SH allegation shall be reported immediately through the Supplier’s and the Purchaser’s grievance channels;

(4) no use of alcohol or prohibited substances on site; no child labour;

(5) follow hazardous-materials, waste-management and spill-response procedures at all times;

(6) respect security instructions and movement protocols applicable to the site; and

(7) protect equipment, samples and data; no removal of the Purchaser’s property from site.

Failure to comply with this Code of Conduct may result in removal of the person from site and other remedies under the Contract.

Acceptance. Acceptance of the Goods requires satisfactory delivery, installation, commissioning, calibration/verification and IQ/OQ/PQ (where applicable), safety checks, cybersecurity configuration and backup/restore testing for networked components, and completion of training and handover documentation, as evidenced by the Operational Acceptance Certificate(s) issued by the Purchaser. The Supplier shall provide test protocols for the Purchaser’s review and approval before testing and shall bear the cost of tests required under the Contract, including retesting required because of Supplier-attributable failure.

Delivery. Bidders may offer delivery earlier than the latest dates stated in the List of Goods and Delivery Schedule; offered delivery dates shall be binding. Bids offering delivery beyond the latest dates stated in Section VII shall be nonresponsive.

A. Analytical Chemistry Equipment

| Item No. | Item | Technical Specification | Qty | Bidder’s Response (Comply / offered specification, with datasheet reference) |
| --- | --- | --- | --- | --- |
| 1 | LC-MS/MS System (Triple Quadrupole) | ESI/APCI source; MRM mode; scan range 50–2000 Da; UHPLC-compatible pump (up to 1200 bar); 100+ vial autosampler; instrument control and data acquisition software; validated for pesticide residues, veterinary drugs, mycotoxins | 1 |  |
| 2 | GC-MS/MS System (Triple Quadrupole) | Triple quadrupole GC-MS/MS; EI source; MRM mode; split/splitless and PTV injector; electronic pressure control; oven ramp 40–450°C; 150+ position autosampler; NIST/Wiley library; instrument control and data acquisition software; validated for pesticide residues, PAHs and contaminants analysis | 1 |  |
| 3 | HPLC System (Quaternary Gradient, DAD + FLD) | Quaternary gradient pump; 0–600 bar; 100-vial autosampler; column oven 10–80°C; DAD 190–800 nm + FLD detector; inline degasser; data acquisition software; MS-ready interface stub | 1 |  |
| 4 | GC-MS System (Single Quadrupole + FID/ECD) | Split/splitless injector; electronic pressure control; temp ramp 40–450°C; quadrupole MS; FID + ECD detectors; headspace sampler; data acquisition software | 1 |  |
| 5 | AAS – Flame + Graphite Furnace (Zeeman) | Flame + graphite furnace mode; Zeeman background correction; hollow cathode lamps for Pb, Cd, As, Hg, Cu, Zn, Fe, Mn; hydride generator; LOD <0.1 ppb (GF mode) | 1 |  |
| 6 | FTIR Spectrometer (Benchtop ATR) | 400–4000 cm⁻¹; resolution ≤0.5 cm⁻¹; DTGS detector; diamond ATR crystal; spectral analysis software; adulteration and species identification library | 1 |  |
| 7 | UV/Vis Spectrophotometer (Double-Beam, Scanning) | 190–1100 nm; spectral bandwidth 0.5–5 nm; stray light <0.05%T; scanning + kinetics mode; flow-through sipper; GMP-compliant software | 1 |  |
| 8 | NIR Analyzer (Rapid Compositional, Benchtop) | 850–2500 nm; sample cup 200 g; pre-loaded calibrations for meat, fish, grains; non-destructive measurement | 1 |  |
| 9 | Ion Chromatograph (Suppressed Conductivity) | Suppressed conductivity detection; anion/cation columns; 25-position autosampler; for nitrate, nitrite, chloride, phosphate, sulfate in food and water matrices | 1 |  |
| 10 | Karl Fischer Titrator (Coulometric) | Coulometric KF; range 10 µg–200 mg H₂O; automatic drift correction; GLP-compliant software; suitable for fats, spices, dried meat | 1 |  |
| 11 | HPLC Column – C18 Reverse Phase (150×4.6 mm, 5 µm) | C18 bonded silica; 150×4.6 mm; 5 µm particle size; for veterinary drugs, vitamins, general HPLC; USP L1 classification | 6 |  |
| 12 | HPLC Column – C8 (150×4.6 mm, 5 µm) | C8 bonded silica; 150×4.6 mm; for lipophilic analytes, beta-agonists, hormone residues; USP L7 | 4 |  |
| 13 | HPLC Column – HILIC (100×2.1 mm, 1.8 µm) | Hydrophilic interaction chromatography; polar analytes (aminoglycosides, mycotoxins, carbohydrates); LC-MS compatible; 100×2.1 mm | 4 |  |
| 14 | HPLC Column – Phenyl-Hexyl (150×4.6 mm, 5 µm) | Enhanced selectivity for polycyclic/aromatic analytes (PAHs, aflatoxins, ochratoxin); USP L11; 150×4.6 mm | 4 |  |
| 15 | HPLC Guard Columns (C18, 12.5×4.6 mm) – Pack of 3 | Matching chemistry to analytical column; 3-pack; protects analytical column; replace monthly under heavy use | 12 |  |
| 16 | GC Capillary Column – DB-5 equivalent (30 m × 0.25 mm × 0.25 µm) | 5% phenyl polysiloxane phase; 30 m; for pesticide residues, PAHs, VOCs; low-bleed coating; bleed <0.5 pA at 300°C | 4 |  |
| 17 | GC Capillary Column – Mid-polarity (30 m × 0.25 mm × 0.25 µm) | 50% phenyl polysiloxane; confirmatory second column; PCBs, organochlorine pesticides | 2 |  |
| 18 | GC Capillary Column – Wax-phase (30 m × 0.25 mm × 0.25 µm) | Polyethylene glycol stationary phase; for fatty acid methyl esters (FAME), aroma compounds, food authenticity | 2 |  |
| 19 | SPE Cartridges – C18 (200 mg / 3 mL), Box 50 | C18 sorbent; 200 mg per cartridge; for matrix cleanup of veterinary drug extracts, mycotoxins; 50 per box | 20 |  |
| 20 | SPE Cartridges – Mixed-mode Dispersive (QuEChERS), Box 50 | Mixed PSA/C18/MgSO4 for QuEChERS dispersive SPE; 50 per box; EN 15662 compliant | 20 |  |
| 21 | Immunoaffinity Columns – Aflatoxin (25/box) | Aflatoxin B1/B2/G1/G2 cleanup; HPLC-FLD compatible; 25 per box; AOAC validated method | 8 |  |
| 22 | Immunoaffinity Columns – Ochratoxin A (25/box) | Ochratoxin A cleanup; 25 per box; AOAC 2000.09 | 6 |  |
| 23 | Immunoaffinity Columns – Fumonisin B1/B2 (25/box) | Fumonisin B1/B2 cleanup; 25 per box; AOAC 2001.06 | 6 |  |

B. Microbiology Equipment

| Item No. | Item | Technical Specification | Qty | Bidder’s Response (Comply / offered specification, with datasheet reference) |
| --- | --- | --- | --- | --- |
| 24 | Biosafety Cabinet – Class II A2 (1.2 m) | Inflow 0.51 m/s; 70% recirculation / 30% exhaust; HEPA H14 supply and exhaust filters; UV lamp interlock; NSF-49 / EN 12469 certified; field-certified on installation | 4 |  |
| 25 | Autoclave / Steam Sterilizer (100 L, Pre-Vacuum) | 134°C / 103 kPa; 100 L chamber; pre-vacuum cycle; integrated data logger; bioseal door; USB printout; EN 13060 | 1 |  |
| 26 | Anaerobic Incubator / Workstation | N₂/CO₂/H₂ atmosphere; O₂ <1 ppm; HEPA filtration; 30–100 L working volume; for Clostridium and anaerobic pathogens | 1 |  |
| 27 | Incubator – Temperature Controlled (100–400 L) | Forced-air convection; 4°C above ambient to 80°C; digital PID controller; ±0.1°C uniformity; 100–400 L; HEPA-filtered option | 4 |  |
| 28 | Real-Time PCR System (96-well, 4–6 channel) | 96-well block; melt-curve analysis; FAM/HEX/ROX/Cy5 channels; pathogen detection panel; ±0.1°C block uniformity; GLP-compliant software | 1 |  |
| 29 | Deep Freezer –80°C (Ultra-Low, 400 L) | -40°C to -86°C; CFC-free refrigerant; temperature alarm; HEPA air filter; 400 L; for culture stocks, enzymes, reference material storage | 2 |  |
| 30 | Refrigerator (2–8°C, 250 L, Lab-Grade) | 2–8°C; ±0.5°C stability; 250 L upright; temperature alarm; CFC-free; explosion-proof option available | 4 |  |
| 31 | Automated Microplate Reader (ELISA) | UV/Vis 200–900 nm; dual-beam; 96/384-well capacity; kinetics mode; for species ID ELISA, antibiotic residues, toxin assays | 3 |  |
| 32 | Colony Counter (Digital, Image Analysis) | LED illuminated plate; pressure-sensitive pad; CFU range 0–9999; LIMS-compatible barcode output; automated image counting | 5 |  |
| 33 | Fluorescence Microscope (Upright) | Plan-apochromat objectives 4×/10×/40×/100× oil; mercury/LED illumination; FITC/TRITC/DAPI filter sets; 12 MP digital camera | 1 |  |
| 34 | Impedance / ATP Microbial Analyzer | ATP luminometry for hygiene surface swabs; 10 RLU sensitivity; USB data output; cloud logging capability; for environmental monitoring | 1 |  |
| 35 | Culture Media – Annual Supply (Selective + Non-selective, 14 types) | Plate Count Agar, VRBA, ALOA, XLD+RVS+MKTTn, mCCDA, CT-SMAC, Baird-Parker, TCBS, DG18, Buffered Peptone Water, MRD, BHI, SDA; 500 g dehydrated per type | 150 |  |
| 36 | Reference Cultures – Freeze-dried (8 species, certified) | E. coli ATCC 25922 equivalent, Salmonella, Listeria, S. aureus, Campylobacter, Vibrio parahaemolyticus, C. perfringens, P. aeruginosa; internationally certified strains | 1 |  |

C. Sample Preparation and General Laboratory Equipment

| Item No. | Item | Technical Specification | Qty | Bidder’s Response (Comply / offered specification, with datasheet reference) |
| --- | --- | --- | --- | --- |
| 37 | Analytical Balance (High-Precision, 220 g / 0.01 mg) | Dual-range 120/220 g; 0.01 mg readability; internal calibration; draft shield; ionizer; RS-232/USB/Ethernet; GLP printout; IP54 protection class | 2 |  |
| 38 | Precision Balance (50 kg / 1 g) | High-capacity bench balance for bulk samples; 1 g readability; RS-232/USB; anti-static pan; 400×500 mm platform | 2 |  |
| 39 | pH Meter (Bench-top, Multi-parameter, IP54) | 0–14 pH; ±0.002 pH accuracy; automatic temperature compensation; dual-junction food-grade electrode; multi-point calibration; GLP logs; IP54 | 4 |  |
| 40 | Moisture Analyzer (Halogen/IR) | Capacity 120 g; 0.001 g readability; temperature 50–230°C; programmable drying profiles; metal pan set; suitable for meat/grains/powders | 3 |  |
| 41 | Refractometer (Digital, Peltier-controlled) | Brix 0–95%; RI 1.3200–1.5300; ±0.00002 RI; Peltier temperature control ±0.05°C; 0.3 mL sample volume; ICUMSA/ISO compliant | 2 |  |
| 42 | Homogenizer / Stomacher – 400 mL (Paddle Blender) | Paddle blender; 80–400 mL sample bags; 10–400 rpm; ISO 6887 compliant; spill-proof seal | 2 |  |
| 43 | Stomacher – High-Volume (3500 mL) | 400–3500 mL capacity bags; variable speed; timer; ISO 6887; for large fish and meat sample preparation | 1 |  |
| 44 | High-Speed Centrifuge – Refrigerated (0–25,000 rpm) | Swing-bucket + fixed-angle rotors; 15/50/100 mL tubes; 4°C operating temperature; for fat separation, serum preparation | 1 |  |
| 45 | Microcentrifuge (Benchtop, 0–16,000 rpm) | 24 × 1.5/2.0 mL tube capacity; for PCR and molecular sample preparation | 2 |  |
| 46 | Water Bath / Heating Bath (5–99°C) | Stainless steel tank; ±0.1°C temperature stability; 10 L capacity; shaking or static; for enzyme assays, sample warm-holding | 3 |  |
| 47 | Dry Bath / Block Heater (Ambient+5 to 130°C) | Interchangeable blocks for 0.5/1.5/2.0 mL and 15/50 mL tubes; ±0.5°C accuracy | 3 |  |
| 48 | Orbital Shaker / Incubator Shaker | 25–500 rpm; temperature range 4–60°C; flask clamps 50 mL–5 L; enzyme assays, microbiology enrichments | 2 |  |
| 49 | Vortex Mixer (Variable 100–3200 rpm) | Continuous and touch mode; multi-tube adapters included | 4 |  |
| 50 | Magnetic Stirrer / Hot Plate (Combined) | Ceramic top surface; 50–380°C; 50–1500 rpm; external temperature probe; closed-loop temperature control | 6 |  |
| 51 | Vacuum Oven (100 L, Ambient–250°C) | 100 L chamber; vacuum <0.1 mbar; programmable ramp/soak profiles; chemical-resistant inner chamber | 1 |  |
| 52 | Muffle Furnace – Ash Determination (1200°C max) | PID controller; ramp/soak temperature profiles; chimney port; 5 L chamber; crucible and porcelain accessories | 1 |  |
| 53 | Drying Oven – Forced Air (Ambient to 300°C) | ±1°C temperature uniformity; forced-air convection; 108 L; glassware sterilization and moisture determination | 3 |  |
| 54 | Ultrasonic Cleaner (20 L, 40 kHz) | Sweep frequency; digital timer and temperature; degas mode; stainless steel tank; for glassware and instrument part cleaning | 2 |  |
| 55 | Water Purification System (Type I + Type II) | Type I: 18.2 MΩ·cm; TOC <5 ppb; UV sterilization + 0.22 µm filtration; compliant with CLSI and ISO 3696 Type I water standards. Type II: RO+UV; 10–20 L/h general lab use | 1 |  |
| 56 | Turbidimeter / Nephelometer (0–4000 NTU) | ±2% full scale; EPA 180.1 compliant; stray-light corrected; for seafood processing water quality | 3 |  |
| 57 | Dissolved Oxygen / Multi-parameter Meter | DO 0–200%/20 mg/L; pH, conductivity, TDS, salinity, temperature; automatic temperature compensation; 12-month cap maintenance kit | 2 |  |
| 58 | Water Activity Meter (0–1.00 aw) | ±0.003 accuracy; automatic temperature compensation; 0.5 mL sample; 5-min measurement; calibration standards included | 2 |  |
| 59 | Texture Analyzer (0.01–250 N) | Load cells 1–50 N and 250 N; shear probe set; compression cell; 0.01–40 mm/s crosshead speed | 1 |  |
| 60 | Colorimeter (Meat/Food Colour) | SCI/SCE measurement; d/8° geometry; CIELab colour space; ΔE <0.05; for meat colour classification per EU/USDA standards | 1 |  |
| 61 | Pipette Set – Single Channel (6 volume ranges) | 0.1–2.5 µL, 0.5–10 µL, 2–20 µL, 10–100 µL, 20–200 µL, 100–1000 µL; autoclavable; GLP-ready; ISO 8655 compliant | 6 |  |
| 62 | Multi-Channel Pipette (8 & 12-channel, 50–300 µL) | Adjustable tip spacing; for ELISA plate loading; 8-channel and 12-channel models | 4 |  |

D. Infrastructure and Support Equipment

| Item No. | Item | Technical Specification | Qty | Bidder’s Response (Comply / offered specification, with datasheet reference) |
| --- | --- | --- | --- | --- |
| 63 | UPS System – Online, 10 kVA (for LC-MS/MS, HPLC, AAS) | Online double-conversion; 10 kVA; ±1% output voltage; 30-min battery backup at full load; remote monitoring port | 4 |  |
| 64 | UPS System – Online, 3 kVA (for PCR, Freezers, Incubators) | Online double-conversion; 3 kVA; ±1% voltage regulation; 20-min backup; for biological storage and PCR systems | 1 |  |
| 65 | Power Line Conditioner / Voltage Stabilizer (5 kVA) | Servo-controlled; ±1% output; 5 kVA; for analytical instruments on shared circuits | 6 |  |
| 66 | HVAC – Dedicated Lab Precision Air Handling Unit | Split unit or central AHU; temperature 18–26°C ±1°C; RH 40–60% ±5%; HEPA pre-filter; for analytical chemistry lab | 2 |  |
| 67 | Fume Hood – Ducted Chemical (1.2–1.5 m) | Face velocity 0.5 m/s at 450 mm sash; chemical-resistant liner; integral spill tray; LED lighting; ASHRAE 110 performance standard | 1 |  |
| 68 | Sample Reception Area Setup | Stainless steel reception bench 1.8 m; waterproof work surface; sample log book station; printed barcode label station; chain-of-custody forms | 1 |  |
| 69 | Cold Storage Tracking System (Integrated with LIMS) | Wall-mounted temperature log display; sample location map display; compatible with LIMS barcode scan in/out | 1 |  |
| 70 | Sample Storage Racks – Traceable (Cryogenic + Ambient) | Cryogenic box racks for -80°C storage (5×5 grid, 81 positions); ambient sample racks (stainless steel, numbered); QR/barcode labeled | 20 |  |
| 71 | Refrigerator -20°C (Intermediate Sample Storage, 200 L) | Upright; -20°C ±2°C; alarm; for extracted samples, PCR reagents, intermediate storage | 2 |  |
| 72 | Cold Chain Maintenance (Portable Coolers, PCM Ice Packs) | Portable cooler boxes with phase-change ice packs; 2–8°C hold ≥24 h; for field collection and transit | 10 |  |
| 73 | Data Logger – Temperature/Humidity (USB, Alarm) | -40 to +85°C; ±0.3°C accuracy; RH 0–100%; traceable calibration certificate | 10 |  |
| 74 | Emergency Shower & Eyewash Station – Plumbed (ANSI Z358.1) | Combination unit: tepid water; dual-nozzle eyewash + drench shower; stainless steel bowl; wall-mounted; ANSI Z358.1 compliant | 3 |  |
| 75 | Portable Eyewash Station (Self-contained, gravity-fed) | 15-min flushing capacity; ANSI Z358.1 compliant; for areas without plumbed emergency access | 4 |  |
| 76 | Chemical Spill Kit (Acid + Base + Solvent) | Neutralizing agents; absorbent pads; disposal bags; gloves; goggles; 20 L capacity each | 8 |  |
| 77 | Biohazard Waste Containers (Autoclavable Bags + Sharps) | Biohazard autoclave bags 20 L and 50 L; sharps containers 5 L; for microbiological lab waste | 20 |  |
| 78 | Chemical Waste Storage Cabinet (Acid-Resistant, 90 L) | Welded steel; acid/base/solvent resistant; self-closing doors; local ventilation port; 90 L capacity | 4 |  |
| 79 | CO₂ Fire Extinguisher (5 kg, EN 3) | For electrical and chemical fires; wall-mounted bracket; annual service | 10 |  |
| 80 | Photoelectric Smoke Detector (IP44, Lab-grade) | Connected alarm; for chemical vapour lab areas; IP44 rated | 8 |  |
| 81 | First Aid Kit (Laboratory Grade, 50-person) | Chemical burns, cuts; sterile dressings; burn gel; eye wash; wall-mountable | 4 |  |

E. Enabling Systems (IT, LIMS and Calibration Support Equipment)

| Item No. | Item | Technical Specification | Qty | Bidder’s Response (Comply / offered specification, with datasheet reference) |
| --- | --- | --- | --- | --- |
| 82 | Barcode Label Printer (Lab-grade, 300 dpi, Thermal) | Direct thermal; 300 dpi resolution; prints GS1-compatible barcodes and QR codes; LIMS-integrated; label sizes 12–100 mm | 4 |  |
| 83 | Barcode Scanner (USB + Wireless, 1D/2D) | Reads GS1-128, Data Matrix, QR Code; USB and Bluetooth connectivity; IP42 rated; for LIMS sample scan-in/out | 8 |  |
| 84 | Role-Based Access Control System (LIMS + Instrument Access) | RFID card reader hardware (6 units); access control software module within LIMS; audit log for all instrument logins; ISO 17025 §6.2 competence tracking | 1 |  |
| 85 | Environmental Monitoring Sensors (Wireless, Temp/RH) | -40 to +85°C; ±0.3°C accuracy; RH 0–100%; replaceable batteries; mesh network compatible | 20 |  |
| 86 | Certified Weight Set – OIML Class E2/F1 (1 mg to 2 kg) | Stainless steel; OIML Class E2/F1; traceable to national standard; for in-house balance calibration; includes storage case | 3 |  |
| 87 | Pipette Calibration System – Gravimetric (ISO 8655) | Calibration balance (0.1 mg); ISO 8655 calculation software; draft shield; evaporation trap; covers all volume ranges 0.1 µL–10 mL | 1 |  |
| 88 | Temperature Calibration – Dry Block Calibrator + Reference Thermometer | Dry block calibrator 0–300°C; ±0.05°C accuracy; interchangeable inserts; traceable Pt100 reference probe; for oven, incubator, autoclave in-house calibration | 1 |  |
| 89 | Calibration Thermometer – Handheld, 4-channel, Pt100 | 4-channel; 0.01°C resolution; -200 to +600°C range; traceable certificate | 4 |  |
| 90 | Laboratory Server / NAS Backup System | Rackmount NAS server; 20 TB RAID-5 configuration; automatic nightly backup to cloud (encrypted); LIMS data host; role-based access; UPS-protected | 1 |  |
| 91 | Network Infrastructure (Managed Switch + Lab Wi-Fi) | Managed PoE switch (24-port); dual-band Wi-Fi access points (3 units); for LIMS, instrument data transfer, environmental monitoring system | 1 |  |
| 92 | Analyst Workstation – Instrument-Dedicated (Desktop) | 64-bit octa-core processor (latest generation); 32 GB RAM; 1 TB NVMe SSD; 24" FHD monitor; 64-bit operating system; wired Ethernet; USB 3.0 hub; pre-imaged for instrument control software | 4 |  |
| 93 | Lab Analyst PC – LIMS & Reporting (Desktop) | 64-bit quad-core processor; 16 GB RAM; 512 GB SSD; 24" FHD monitor; 64-bit operating system; LIMS client; for data entry, COC processing, report drafting and results review | 6 |  |
| 94 | Lab Manager / QA Officer Laptop | 64-bit quad-core processor; 16 GB RAM; 512 GB SSD; 15.6" FHD display; 64-bit operating system; encrypted hard drive; for management, ISO 17025 document control, audit trail review, external reporting | 2 |  |
| 95 | Network Laser Printer – A4/A3 (Mono + Colour, Duplex) | A4/A3 capable; mono and colour; automatic duplex; network (Ethernet + Wi-Fi); 30 ppm mono; for test reports, calibration certificates, COC forms | 2 |  |
| 96 | UPS – Line-Interactive (1 kVA) for Workstations & PCs | Line-interactive; 1 kVA / 600 W; ±10% voltage regulation; 10-min backup at full load; surge-protected outlets ×6; USB monitoring port | 12 |  |
| 97 | External Hard Drive – Encrypted Backup (4 TB, USB 3.0) | Hardware AES 256-bit encryption; USB 3.0; 4 TB; bus-powered; for instrument data and LIMS report backups | 6 |  |

F. Livestock and Feed Testing Equipment

| Item No. | Item | Technical Specification | Qty | Bidder’s Response (Comply / offered specification, with datasheet reference) |
| --- | --- | --- | --- | --- |
| 98 | Milk Analyzer – Multi-parameter (Fat/Protein/Lactose/SNF/Conductivity) | Ultrasonic + infrared measurement principle; fat, protein, lactose, SNF, density, conductivity, freezing point; <60 s per sample; IDF/ISO 9622 compliant; 200+ samples/day throughput | 1 |  |
| 99 | Hematology Analyzer – Automated CBC (Livestock Blood Panels) | Automated 5-part differential; WBC, RBC, HGB, HCT, PLT, MCV, MCH, MCHC; for cattle, sheep, goat, camel blood panels; 60 samples/hour; EDTA whole blood | 1 |  |
| 100 | Veterinary Biochemistry Analyzer (Photometric, Automated) | Photometric; 200–700 nm; 200 tests/hour; for glucose, urea, creatinine, ALT, AST, total protein, albumin, electrolytes (Na/K/Ca/Cl) in serum/plasma | 1 |  |
| 101 | Portable Milk Tester – Field Grade | Ultrasonic; battery-powered; measures fat, SNF, protein, density, added water; <30 s; suitable for farm-level screening | 1 |  |
| 102 | Gerber Centrifuge (12-position, ISO 2446) | 12-position; 1000–1500 rpm; for milk fat by Gerber/Babcock reference method; ISO 2446; temperature-controlled chamber | 1 |  |
| 103 | Portable Livestock Disease Rapid Test Reader | Lateral flow reader; reads 20 strip formats; semi-quantitative; LIMS-integrated via USB/Bluetooth; for on-site disease screening | 1 |  |
| 104 | Stereomicroscope (Feed Adulteration Examination) | 10–40× magnification; transmitted and incident illumination; zoom objective; for bone meal, feather meal, blood meal, insect fragment identification in compound feed | 2 |  |
| 105 | Feed Particle Sieve Analyser (Vibrating Sieve Tower) | Vibrating sieve tower; 5-screen set (0.25–4.75 mm); timer; amplitude control; for pellet/grind size distribution per ISO 6497 | 1 |  |
| 106 | ELISA Kit – Brucellosis (B. abortus + B. melitensis) | Plate-based competitive ELISA; OIE validated; 96-well; for cattle, small ruminant serum; includes positive/negative controls | 5 |  |
| 107 | ELISA Kit – Foot & Mouth Disease (FMD, Serotypes O/A/SAT) | Non-structural protein ELISA; OIE validated; detects all SAT serotypes; 96-well; for cattle, sheep, goat, camel | 5 |  |
| 108 | ELISA Kit – African Swine Fever (ASFV Antigen + Antibody) | Antigen capture + antibody detection; OIE validated; 96-well; for pig serum and tissue homogenate | 10 |  |
| 109 | ELISA Kit – Peste des Petits Ruminants (PPR, cELISA) | Competitive ELISA; OIE validated; for sheep and goat serum; detects IgG against PPR virus | 10 |  |
| 110 | ELISA Kit – Rift Valley Fever (RVF, IgM/IgG) | IgM capture and IgG indirect ELISA; OIE validated; for cattle, sheep, goat, camel serum | 5 |  |
| 111 | ELISA Kit – Bovine Tuberculosis (IFN-gamma, 24-well) | IFN-gamma stimulation and detection; 24-well format; for whole blood from cattle; OIE validated | 10 |  |
| 112 | ELISA Kit – Avian Influenza H5/H7 (Competitive ELISA) | Competitive ELISA; OIE validated; 96-well; for poultry serum; H5/H7 subtype differentiation | 10 |  |
| 113 | Lateral Flow Strip – Beta-Lactam Antibiotics (Milk + Tissue) | Rapid lateral flow; detection at EU MRL; milk and tissue matrices; 50 per box; 3-min result | 4 |  |
| 114 | Lateral Flow Strip – Tetracyclines (4-in-1 card) | Simultaneous OTC/CTC/TC/DC detection; EU MRL level; milk and tissue; 3-min result; 50 per box | 4 |  |
| 115 | Lateral Flow Strip – Chloramphenicol (Banned Substance) | Banned substance detection; urine and tissue; detection <0.3 ppb; 25 per box | 2 |  |
| 116 | Lateral Flow Strip – Streptomycin / Aminoglycosides | Aminoglycoside class detection; EU MRL; 50 per box; milk and meat matrices | 2 |  |
| 117 | Lateral Flow Strip – Clenbuterol (Beta-agonist Growth Promoter) | Banned growth promoter; clenbuterol detection <0.1 ppb; urine and tissue; 25 per box | 4 |  |
| 118 | Lateral Flow Strip – Aflatoxin M1 in Milk | EU limit 0.05 µg/kg; rapid 5-min test; 25 per box; milk matrix only | 4 |  |
| 119 | Blood Collection Tubes – EDTA (10 mL, 1000/box) | EDTA anticoagulant; 10 mL; vacuum; for haematology (CBC) and molecular testing | 5 |  |
| 120 | Blood Collection Tubes – Plain/SST (10 mL, 1000/box) | Serum separator; 10 mL; vacuum; for biochemistry and serology | 5 |  |
| 121 | Vacutainer Needles + Holders (21G × 1.5", 100/box) | 21-gauge × 1.5"; multi-sample needle; holder included; sterile; single-use | 30 |  |
| 122 | Animal Ear Tag Labels + Barcode Stickers (1000/roll) | Livestock traceability; UV-resistant; tamper-evident; barcode-compatible with LIMS | 10 |  |
| 123 | Feed Sampling Probe (Stainless Steel, 1.5 m) | 5-section stainless steel grain/pellet/forage probe; 1.5 m extended; for representative sampling of feed lots per ISO 6497 | 5 |  |
| 124 | Feed CRM – Proximate Analysis (Protein/Fat/Moisture/Ash/Fibre) | ISO 17034 certified; certified values for crude protein, ether extract, moisture, ash, crude fibre; 50 g; suitable for compound feed matrix | 10 |  |
| 125 | Milk CRM – Compositional Analysis (Fat/Protein/Lactose) | ISO 17034 certified; certified fat, protein, lactose, SNF values; freeze-dried; 10 g vial | 6 |  |

G. Consumables and Reagents (First-Year Supply)

| Item No. | Item | Technical Specification | Qty | Bidder’s Response (Comply / offered specification, with datasheet reference) |
| --- | --- | --- | --- | --- |
| 126 | Acetonitrile – HPLC Grade (2.5 L) | ≥99.8% purity; HPLC mobile phase and QuEChERS extraction | 20 |  |
| 127 | Methanol – HPLC Grade (2.5 L) | ≥99.8%; HPLC mobile phase, lipid/drug extraction | 20 |  |
| 128 | Hexane – HPLC Grade (2.5 L) | ≥96%; fat extraction, lipid class separation | 15 |  |
| 129 | Acetone – HPLC Grade (2.5 L) | ≥99.8%; QuEChERS pesticide extraction | 15 |  |
| 130 | Dichloromethane – HPLC Grade (2.5 L) | Pesticide residue extraction, fat cleanup | 10 |  |
| 131 | Ethanol – AR Grade (2.5 L) | 95–96%; general solvent, standard preparation | 15 |  |
| 132 | Acetic Acid – Glacial AR (2.5 L) | ≥99.8%; buffer preparation, mobile phase additive | 10 |  |
| 133 | Nitric Acid – Trace Metal Grade (500 mL) | 65–68%; AAS/ICP sample digestion | 8 |  |
| 134 | Hydrochloric Acid – AR Grade (2.5 L) | 37%; protein hydrolysis, pH adjustment | 10 |  |
| 135 | Sulfuric Acid – AR Grade (2.5 L) | 96–98%; Kjeldahl digestion, fat determination | 10 |  |
| 136 | Sodium Hydroxide – AR Pellets (1 kg) | Kjeldahl distillation, pH adjustment | 10 |  |
| 137 | Heavy Metal Standards – 1000 mg/L (8 elements) | Certified reference solution; 100 mL; for AAS calibration per Codex MRL; elements: Pb, Cd, As, Hg, Cu, Zn, Fe, Mn | 24 |  |
| 138 | Pesticide Reference Standard Mix (100+ compounds) | EU/Codex MRL-level compounds; 10 µg/mL; 1 mL ampoule; organophosphates, pyrethroids, triazines, carbamates; GC/LC-MS grade | 6 |  |
| 139 | Veterinary Drug Standard Mix (50–100 compounds) | Tetracyclines, β-lactams, sulfonamides, aminoglycosides, macrolides; 1 mg/mL stock; LC-MS grade; certified | 6 |  |
| 140 | Mycotoxin Standards (AFB1/B2/G1/G2, OTA, Fumonisin B1/B2) | Certified reference material; 100 µg/mL; ISO 17034 certified; 1 mL vial each | 12 |  |
| 141 | Certified Reference Material – Meat/Fish Matrix (ISO 17034) | Freeze-dried CRM for heavy metals, fat, protein, moisture; ISO 17034 certified; 5 g per vial | 10 |  |
| 142 | pH Buffer Standards (pH 4.00 / 7.00 / 10.00, NIST-traceable) | ±0.01 pH accuracy at 25°C; 250 mL; traceable to international standards; for daily pH meter calibration | 24 |  |
| 143 | Karl Fischer Reagent (Composite grade, 1 L) | Titer 5 mg H₂O/mL; for moisture in fats, spices, dried meat | 10 |  |
| 144 | Argon Gas – Purity ≥99.995% (8500 L cylinder) | AAS/GF-AAS purge and nebulizer gas; with regulator | 4 |  |
| 145 | Nitrogen Gas – Purity ≥99.999% (8500 L cylinder) | GC carrier gas; LC-MS source gas; evaporation; with regulator | 4 |  |
| 146 | Acetylene Gas – Dissolved ≥99.5% (8 kg cylinder) | AAS flame fuel; with flashback arrestor and regulator | 4 |  |
| 147 | Borosilicate Glassware Sets (Beakers, Flasks, Cylinders, Volumetric) | Class A volumetric; borosilicate 3.3; ISO 3585; includes 50–2000 mL beakers, flasks, cylinders | 10 |  |
| 148 | Centrifuge Tubes (50 mL + 15 mL PP), DNase/RNase-free | 500 tubes/case; sterile; self-standing | 20 |  |
| 149 | Microcentrifuge Tubes (1.5 mL + 2.0 mL), PCR-clean | 1000/box; DNase/RNase-free | 30 |  |
| 150 | PCR Plates (96-well, 0.2 mL, thin-wall) | Clear/white; 25 plates/box; certified for qPCR | 30 |  |
| 151 | Pipette Tips (10 µL, 200 µL, 1000 µL) – Filtered & Unfiltered | Autoclavable; filtered and unfiltered; 96-tip rack/box; DNase/RNase-free | 60 |  |
| 152 | Filter Papers – Grades 1, 4, 41, 42 (100/box) | Qualitative and quantitative; 90 mm and 125 mm; cellulose | 20 |  |
| 153 | Membrane Filters – 0.2 µm & 0.45 µm (47 mm, 50/box) | Sterile; individually wrapped; HPLC filtration and microbiological water testing | 20 |  |
| 154 | Disposable Nitrile Gloves – Powder-free (100/box, S/M/L) | 10-mil; extended cuff; CE/FDA certified | 200 |  |
| 155 | Lab Coats – White, Knee-length | CE EN 13034; reinforced stitching; S/M/L sizes | 60 |  |
| 156 | FFP2 / N95 Masks (Non-valve, EN 149/NIOSH) | For mycotoxin and dry powder handling; 200/box | 200 |  |
| 157 | Fat Extraction Thimbles – Cellulose (33×80 mm, 100/box) | Defatted; standardized dimensions; for automated fat extraction | 20 |  |
| 158 | Rapid Test Kits – Livestock Disease Panel (Annual, 6 types) | ASFV, Brucellosis, BTB, Avian Influenza, FMD, Newcastle Disease; plate-based ELISA; OIE validated; 8–20 kits per type | 80 |  |
| 159 | Rapid Test Kits – Meat & Seafood Safety (Annual, 8 types) | Antibiotic residues, histamine, clenbuterol, PSP, DSP, aquaculture antibiotics, meat speciation, multi-residue; 10–40 kits per type | 100 |  |

H. Technical Requirements for Related Services

H.1 Installation, Commissioning and Qualification (IQ/OQ/PQ).

All instruments shall be installed and commissioned at the SoBS National Reference Laboratory, Mogadishu, by factory-certified engineers of the respective manufacturers.

The Supplier shall:

(a) verify site readiness and utility requirements (power, gases, ventilation, water, drainage, network) and advise the Purchaser of any deficiencies before delivery;

(b) perform Installation Qualification (IQ), Operational Qualification (OQ) and Performance Qualification (PQ) in accordance with the manufacturer’s documented protocols for each applicable instrument;

(c) furnish calibration certificates traceable to national or international standards at hand-over; and

(d) sign off jointly with the Purchaser a formal acceptance certificate for each instrument upon successful completion of the agreed acceptance criteria.

H.2 Training.

On-site training courses shall be delivered in English and Somali as listed in the List of Related Services and Completion Schedule (HPLC, GC-MS, AAS and qPCR instrument training, and LIMS administrator and user training). Each course shall include a curriculum agreed in advance, comprehensive training materials handed over to the Purchaser, hands-on operation by trainees, and a competency assessment with certificates of completion. Unless otherwise stated, each instrument course shall accommodate four to six (4–6) SoBS analysts; LIMS training shall cover two (2) system administrators and ten (10) end-users.

H.3 Proficiency Testing (PT) Enrolment.

Enrolment of the NRL with an ISO/IEC 17043-accredited proficiency testing provider for ten (10) rounds in the first year of operation, covering as a minimum: mycotoxins, pesticide residues, heavy metals, nutritional composition, and microbiology schemes.

H.4 External Calibration Services.

Fifteen (15) calibration events in the first year of operation, performed by an ISO/IEC 17025-accredited calibration laboratory, covering balances, instrument qualification support (HPLC/GC), temperature mapping of refrigerators, freezers, incubators and ovens, and pipette verification, with certificates issued for each event.

H.5 Warranty and After-Sales Support.

A minimum warranty of twenty-four (24) months from the date of Operational Acceptance of each instrument (or thirty (30) months from delivery, whichever occurs first), covering parts and labour; remote diagnostic support with acknowledgement and remote troubleshooting within two (2) Business Days of notification; on-site attendance within five (5) Business Days for critical equipment and ten (10) Business Days for non-critical equipment, where local or regional service support is required and site access and security conditions permit; restoration or an agreed corrective action plan within thirty (30) calendar days; at least two (2) preventive maintenance visits per year during the warranty period; loan or replacement equipment and an escalation plan for downtime exceeding thirty (30) consecutive days on critical equipment; and assured availability of spare parts and consumables for at least five (5) years from Operational Acceptance.

4. Drawings

This bidding document includes No Drawings

| List of Drawings |  |  |
| --- | --- | --- |
| Drawing Nr. | Drawing Name | Purpose |

5. Inspections and Tests

The following inspections and tests shall be performed:

(a) Pre-shipment verification: the Supplier shall provide, prior to shipment, manufacturer’s certificates of conformity, test reports and calibration certificates (where applicable) evidencing that each item meets the specifications of Section VII; the Purchaser or its appointed agent may, at its option and cost, carry out pre-shipment inspection at the manufacturer’s premises.

(b) Post-delivery inspection, testing and acceptance: all Goods will be inspected on arrival at the final destination (SoBS – NRL, Mogadishu) for quantity, completeness and transit damage. Final acceptance testing shall be conducted at the final destination after installation and commissioning, at the Supplier’s cost, and shall comprise the successful completion of the agreed IQ/OQ/PQ protocols and acceptance criteria for each instrument, witnessed by the Purchaser. Each instrument will be accepted only upon signature of its acceptance certificate; any item failing to meet the required specifications shall be replaced or remedied at the Supplier’s cost.

ANNEX VII-A TO SECTION VIIS

CONTRACTOR’S ENVIRONMENTAL AND SOCIAL MANAGEMENT PLAN (C-ESMP): MINIMUM REQUIREMENTS

Note on application: This Annex sets out the minimum requirements for the Contractor’s Environmental and Social Management Plan (C-ESMP) which the Supplier shall prepare, submit and implement in accordance with the Special Conditions of Contract (SCC GCC 14.10). In this Annex, “Contractor” means the Supplier under the Contract, “PIU” means the Purchaser, and “Contract” means the contract resulting from this bidding document. Placeholders and signature blocks in this Annex are to be completed at contract implementation stage; nothing in this Annex is required to be completed or submitted at bid stage, except the signed ESHS/OHS commitment statement required under Section III.

1. INTRODUCTION

The Government of Somalia, through the Ministry of Finance Project Implementation Unit (PIU), is supporting the supply, delivery, installation, testing and commissioning of equipment for the National Reference Laboratory (NRL).

Although the activity does not involve major civil works, it may create environmental, social, health and safety risks during equipment transportation, unloading, temporary storage, installation, electrical connection, testing, commissioning and handover.

This Contractor’s Environmental and Social Management Plan (C-ESMP) establishes the measures that the Contractor should implement to prevent, minimize and manage these risks.

The Contractor shall implement this C-ESMP in accordance with the Contract, applicable Somali requirements, relevant World Bank Environmental and Social Standards, and the DRIVE Project environmental and social instruments.

The Contractor shall not commence installation activities until the C-ESMP has been reviewed and found acceptable by the PIU and, where required, the World Bank.

2. OBJECTIVES AND SCOPE

The objectives of the C-ESMP are to:

•  identify E&S and occupational health and safety risks;

•  establish practical mitigation measures;

•  protect workers, NRL staff, visitors and nearby communities;

•  ensure safe transportation, unloading and equipment installation;

•  prevent environmental pollution and poor waste management;

•  manage electronic waste and damaged equipment appropriately;

•  prevent SEA/SH and other worker misconduct;

•  provide accessible grievance mechanisms;

•  establish emergency preparedness and incident reporting procedures; and

•  establish clear monitoring and reporting responsibilities.

The C-ESMP applies to all Contractor workers, subcontractors, suppliers and other persons engaged in:

•  transportation and delivery;

•  loading and unloading;

•  temporary storage;

•  unpacking and positioning;

•  installation and fixing;

•  electrical and other utility connections;

•  testing, calibration and commissioning;

•  waste and packaging management;

•  training;

•  final cleaning and handover.

3. APPLICABLE ENVIRONMENTAL AND SOCIAL REQUIREMENTS

The Contractor shall comply with applicable requirements of:

•  the World Bank Environmental and Social Framework;

•  ESS1 – Assessment and Management of Environmental and Social Risks and Impacts;

•  ESS2 – Labor and Working Conditions;

•  ESS3 – Resource Efficiency and Pollution Prevention and Management;

•  ESS4 – Community Health and Safety;

•  ESS10 – Stakeholder Engagement and Information Disclosure;

•  DRIVE Project Environmental and Social Commitment Plan (ESCP);

•  DRIVE Project Environmental and Social Management Framework (ESMF);

•  applicable ESMP requirements;

•  Stakeholder Engagement Plan (SEP);

•  Project Implementation Manual and associated E&S procedures;

•  applicable Somali environmental, labor and OHS requirements; and

•  relevant manufacturer safety instructions and Good International Industry Practice.

The Contractor shall ensure that all subcontractors and suppliers working at the site comply with the same requirements.

4. CONTRACTOR ESHS/OHS COMMITMENT

Before mobilization, the Contractor shall submit a signed ESHS/OHS Commitment Statement confirming that it will:

1.  Prepare and implement a contract-specific C-ESMP acceptable to the Purchaser before mobilization.

2.  Comply with applicable World Bank Environmental and Social Standards and DRIVE Project E&S instruments.

3.  Provide appropriate ESHS/OHS supervision and adequate resources.

4.  Conduct worker induction and task-specific safety briefings.

5.  Provide suitable PPE to all workers.

6.  Implement safe systems of work for unloading, lifting, electrical works, installation, testing and commissioning.

7.  Enforce a Worker Code of Conduct.

8.  Prevent sexual exploitation and abuse and sexual harassment.

9.  Establish a Worker Grievance Mechanism.

10.  Cooperate with the DRIVE Project GRM.

11.  Maintain emergency preparedness and first-aid arrangements.

12.  Notify and report accidents and incidents in accordance with the Contract and SCC.

13.  Monitor implementation of the C-ESMP and address non-compliance promptly.

The statement shall be signed by an authorized representative of the Contractor.

5. ROLES AND RESPONSIBILITIES

5.1 Contractor

The Contractor shall have primary responsibility for implementation of this C-ESMP. It shall:

•  appoint a competent ESHS/OHS focal person;

•  provide resources for ESHS implementation;

•  conduct risk assessments and Job Safety Analyses;

•  provide worker induction and toolbox talks;

•  ensure PPE use;

•  supervise high-risk activities;

•  manage waste and e-waste;

•  maintain grievance and incident records;

•  implement corrective actions; and

•  provide ESHS reports to the PIU.

5.2 Contractor ESHS/OHS Focal Person

The ESHS/OHS focal person shall:

•  conduct routine inspections;

•  monitor PPE and safe work practices;

•  maintain ESHS records;

•  support worker induction;

•  monitor Code of Conduct compliance;

•  record incidents and near misses;

•  monitor waste management;

•  track corrective actions; and

•  coordinate with the PIU E&S Specialist.

5.3 PIU

The PIU shall:

•  review and clear the C-ESMP;

•  monitor Contractor compliance;

•  conduct site inspections;

•  review ESHS reports;

•  monitor corrective actions;

•  provide E&S guidance; and

•  coordinate with the World Bank on serious incidents and significant E&S matters.

5.4 SoBS/NRL

SoBS/NRL shall facilitate safe access to work areas, coordinate site activities and support implementation of emergency and security arrangements.

6. KEY ENVIRONMENTAL AND SOCIAL RISKS AND MITIGATION MEASURES

| Activity/Risk | Potential Impact | Main Mitigation Measures |
| --- | --- | --- |
| Transportation of equipment | Road accidents, injury, equipment damage | Use competent drivers, roadworthy vehicles, secure loads and comply with traffic rules |
| Delivery and unloading | Crush injuries, falling equipment | Controlled unloading zone, trained operators, lifting equipment, banksman and PPE |
| Manual handling | Musculoskeletal injury | Use mechanical aids, trolleys and team lifting where required |
| Temporary storage | Trip hazards, equipment damage | Secure storage, clear access routes and proper stacking |
| Unpacking | Cuts and punctures | Gloves, suitable tools and controlled unpacking |
| Drilling and fixing | Dust, noise, eye injury | Dust control, eye and hearing protection and restricted access |
| Electrical works | Electric shock, fire | Qualified technicians, isolation procedures, grounding, inspection and testing |
| Equipment installation | Cuts, falls, crushing | Task-specific risk assessment, supervision, PPE and housekeeping |
| Work at height | Falls | Appropriate ladders/platforms, inspection and competent workers |
| Testing and commissioning | Electrical/mechanical hazards | Manufacturer procedures, competent staff, controlled access and emergency shutdown |
| Chemicals/reagents | Exposure or spills | Follow SDS and manufacturer instructions; appropriate storage and PPE |
| Packaging waste | Waste accumulation/fire risk | Segregate, reuse, recycle and remove regularly |
| Damaged electronics | E-waste pollution | Secure storage, supplier take-back or approved disposal/recycling |
| Worker misconduct | Harassment, conflict, SEA/SH | Code of Conduct, worker training and disciplinary procedures |
| Fire | Injury or property damage | Fire extinguishers, clear exits and emergency procedures |
| Emergency incidents | Injury/fatality | First aid, trained personnel, evacuation procedures and emergency contacts |

7. OCCUPATIONAL HEALTH AND SAFETY

The Contractor shall provide a safe working environment and apply the hierarchy of controls.

Before high-risk activities, the Contractor shall prepare a Job Safety Analysis identifying:

•  task;

•  hazards;

•  persons exposed;

•  control measures;

•  residual risks; and

•  responsible person.

Minimum PPE shall be determined by task and may include:

•  safety footwear;

•  gloves;

•  work clothing/high-visibility vest;

•  safety glasses;

•  hard hats;

•  hearing protection;

•  dust masks or respiratory protection; and

•  electrical PPE.

Only competent personnel shall perform electrical installation and testing.

Electrical systems should be isolated before work wherever practicable, and circuits should be inspected and tested before energization.

For lifting and unloading, the Contractor shall:

•  determine load weight;

•  inspect lifting equipment;

•  establish exclusion zones;

•  use trained operators and spotters;

•  prohibit workers from standing beneath suspended loads; and

•  stop work where lifting conditions become unsafe.

Good housekeeping shall be maintained throughout the work site.

8. LABOR, CODE OF CONDUCT AND SEA/SH

The Contractor shall comply with ESS2 and applicable labor requirements.

The Contractor shall:

•  prohibit child and forced labor;

•  avoid discrimination;

•  provide safe working conditions;

•  provide access to water, sanitation and first aid;

•  ensure appropriate rest periods; and

•  provide a Worker Grievance Mechanism.

All workers shall sign a Code of Conduct before beginning work.

The Code of Conduct shall prohibit:

•  violence;

•  intimidation;

•  discrimination;

•  sexual exploitation and abuse;

•  sexual harassment;

•  sexual activity with persons under 18;

•  exchange of money, goods or services for sexual favors;

•  alcohol or drug use while working; and

•  other inappropriate conduct.

Workers shall be trained on SEA/SH prevention and reporting.

SEA/SH complaints shall be handled confidentially, sensitively and in accordance with survivor-centered principles. Routine reports should not disclose personally identifying information.

9. GRIEVANCE REDRESS MECHANISMS

9.1 Project/Stakeholder GRM

The Contractor shall cooperate with the DRIVE Project GRM.

Potential complaints may relate to:

•  unsafe work;

•  worker behavior;

•  noise or access restrictions;

•  property damage;

•  waste management;

•  harassment; or

•  other environmental and social concerns.

The Contractor shall inform workers and relevant stakeholders about available grievance channels.

Complaints shall be recorded, referred, assessed, addressed and closed in accordance with Project procedures.

The GRM shall be free, accessible, transparent and free from retaliation.

9.2 Worker GRM

A separate Worker GRM shall allow workers to raise concerns relating to:

•  working conditions;

•  OHS;

•  wages or employment conditions;

•  discrimination;

•  harassment;

•  supervisor behavior; and

•  other workplace issues.

Confidential complaints shall be protected, and no worker shall face retaliation for raising a legitimate concern.

10. WASTE, E-WASTE AND POLLUTION MANAGEMENT

The Contractor shall apply the waste hierarchy:

Avoid – Reduce – Reuse – Recycle – Recover – Dispose Safely.

Waste may include:

•  cardboard;

•  plastic;

•  wooden crates;

•  foam;

•  metal;

•  cables;

•  packaging materials;

•  electronic components;

•  damaged equipment; and

•  minor installation waste.

Waste shall be segregated and stored safely.

Open burning and indiscriminate dumping are prohibited.

Damaged electronic and electrical equipment shall:

•  be identified and segregated;

•  be stored securely;

•  not mixed with municipal waste;

•  not be openly burned or dismantled by unauthorized persons;

•  be returned to the supplier/manufacturer where possible; or

•  be transferred through an environmentally sound disposal or recycling arrangement acceptable under Project requirements.

The Contractor shall also prevent spills, minimize dust and noise and ensure that chemicals, oils or other pollutants are not discharged into soil, drains or water.

11. EMERGENCY PREPAREDNESS, FIRE AND SECURITY

The Contractor shall maintain an Emergency Response Plan covering foreseeable emergencies, including:

•  fire;

•  electrical shock;

•  serious injury;

•  equipment collapse;

•  chemical spill;

•  traffic accident;

•  security incident; and

•  evacuation.

The Contractor shall provide:

•  first-aid kits;

•  trained first-aid personnel;

•  fire extinguishers;

•  emergency contact numbers;

•  evacuation routes;

•  assembly point; and

•  access to medical assistance.

Workers shall be briefed on emergency procedures before commencing work.

The Contractor shall also comply with site security arrangements and coordinate closely with the PIU and SoBS/NRL.

12. INCIDENT AND ACCIDENT REPORTING

The Contractor shall immediately respond to accidents and incidents and take measures to prevent further harm.

Reportable incidents include:

•  fatalities;

•  serious injuries;

•  electrical shock;

•  fire;

•  serious transport accidents;

•  major equipment accidents;

•  significant spills;

•  serious worker misconduct;

•  SEA/SH allegations;

•  security incidents; and

•  significant environmental damage.

The Contractor shall:

1.  provide emergency assistance;

2.  make the area safe;

3.  notify the PIU/Purchaser within the timeframe required by the SCC, Contract and Project procedures;

4.  record the incident;

5.  cooperate with investigation;

6.  identify immediate and root causes;

7.  prepare corrective actions; and

8.  track corrective actions to closure.

Serious incidents shall be reported immediately and shall not be delayed until submission of routine reports.

SEA/SH incidents shall be handled confidentially, and only minimum non-identifying information shall be communicated through standard reporting channels.

13. TRAINING AND AWARENESS

Before starting work, all Contractor personnel shall receive ESHS induction covering:

•  C-ESMP requirements;

•  OHS;

•  PPE;

•  electrical safety;

•  safe lifting/manual handling;

•  emergency response;

•  fire prevention;

•  waste/e-waste management;

•  Code of Conduct;

•  SEA/SH prevention;

•  Worker GRM;

•  Project GRM; and

•  incident reporting.

Toolbox talks should be conducted as necessary before high-risk or unfamiliar activities.

Attendance records shall be maintained.

14. MONITORING AND REPORTING

The Contractor shall monitor implementation of the C-ESMP throughout the installation and commissioning period.

| Parameter | Indicator | Frequency | Responsibility |
| --- | --- | --- | --- |
| Worker induction | Workers trained | Before starting work | Contractor |
| Code of Conduct | Workers signed | Before starting work | Contractor |
| PPE | Compliance | Daily | ESHS Focal Person |
| Electrical safety | Safe installation/testing | Each activity | Qualified technician |
| Lifting operations | Compliance with lifting plan | Each lift | Supervisor |
| Housekeeping | Clean and safe work area | Daily | Contractor |
| Waste management | Proper segregation/removal | Weekly | Contractor |
| E-waste | Appropriate storage/disposal | As generated | Contractor |
| GRM | Complaints received/resolved | Ongoing | Contractor/PIU |
| Incidents | Number and severity | Continuous | Contractor |
| Corrective actions | Closure of actions | Weekly | Contractor/PIU |

The Contractor shall provide periodic ESHS reports to the PIU covering:

•  work undertaken;

•  workforce numbers;

•  induction and toolbox talks;

•  PPE compliance;

•  inspections;

•  incidents and near misses;

•  waste and e-waste;

•  grievances;

•  Code of Conduct compliance;

•  corrective actions; and

•  planned activities.

15. NON-CONFORMANCE AND CORRECTIVE ACTION

Any ESHS non-compliance shall be recorded and corrected promptly.

The Contractor shall:

•  identify the non-conformance;

•  assess its risk;

•  take immediate corrective measures where required;

•  assign responsibility;

•  establish a completion date;

•  document corrective action; and

•  verify closure.

The PIU may require suspension of an activity where unacceptable or imminent ESHS risks have not been adequately controlled.

16. DEMOBILIZATION AND HANDOVER

Before handover, the Contractor shall:

•  remove temporary materials;

•  remove all wastes;

•  clean work areas;

•  appropriately manage damaged electronic components;

•  close outstanding ESHS actions;

•  complete required testing;

•  provide safety and operating documentation;

•  train designated NRL staff; and

•  confirm that the site is safe and environmentally acceptable.

17. CONTRACTOR DECLARATION

The Contractor confirms that it has reviewed and understood this C-ESMP and commits to implementing all applicable environmental, social, health and safety requirements during the supply, delivery, installation, testing and commissioning of National Reference Laboratory equipment.

Contractor: ___________________________

Authorized Representative: ___________________________

Position: ___________________________

Signature: ___________________________

Date: ___________________________

Company Stamp: ___________________________

ANNEXES TO THE C-ESMP

The following should accompany the final C-ESMP:

1.  Signed ESHS/OHS Commitment Statement.

2.  Worker Code of Conduct.

3.  Worker GRM Procedure.

4.  Project GRM Referral Procedure.

5.  SEA/SH Prevention and Response Protocol.

6.  Emergency Response Plan.

7.  Incident Reporting Form.

8.  Waste and E-Waste Register.

9.  Job Safety Analysis Form.

10.  Daily ESHS Inspection Checklist.

11.  Training/Toolbox Talk Register.

12.  PPE Register.

13.  Corrective Action Register.

PART 3 – Contract

Section VIII - General Conditions of Contract

Table of Clauses

1.	Definitions	107

2.	Contract Documents	108

3.	Fraud and Corruption	108

4.	Interpretation	108

5.	Language	109

6.	Joint Venture, Consortium or Association	109

7.	Eligibility	110

8.	Notices	110

9.	Governing Law	110

10.	Settlement of Disputes	110

11.	Inspections and Audit by the Bank	111

12.	Scope of Supply	111

13.	Delivery and Documents	111

14.	Supplier’s Responsibilities	112

15.	Contract Price	113

16.	Terms of Payment	113

17.	Taxes and Duties	114

18.	Performance Security	114

19.	Copyright	114

20.	Confidential Information	115

21.	Subcontracting	116

22.	Specifications and Standards	116

23.	Packing and Documents	116

24.	Insurance	117

25.	Transportation and Incidental Services	117

26.	Inspections and Tests	118

27.	Liquidated Damages	119

28.	Warranty	119

29.	Patent Indemnity	120

30.	Limitation of Liability	121

31.	Change in Laws and Regulations	121

32.	Force Majeure	122

33.	Change Orders and Contract Amendments	122

34.	Extensions of Time	123

35.	Termination	124

36.	Assignment	125

37.	Export Restriction	125

Section VIII. General Conditions of Contract

| Definitions | 1.1 The following words and expressions shall have the meanings hereby assigned to them: ### “Bank” means the World Bank and refers to the International Bank for Reconstruction and Development (IBRD) or the International Development Association (IDA). ### “Contract” means the Contract Agreement entered into between the Purchaser and the Supplier, together with the Contract Documents referred to therein, including all attachments, appendices, and all documents incorporated by reference therein. ### “Contract Documents” means the documents listed in the Contract Agreement, including any amendments thereto. ### “Contract Price” means the price payable to the Supplier as specified in the Contract Agreement, subject to such additions and adjustments thereto or deductions therefrom, as may be made pursuant to the Contract. ### “Day” means calendar day. ### “Completion” means the fulfillment of the Related Services by the Supplier in accordance with the terms and conditions set forth in the Contract. ### “GCC” means the General Conditions of Contract. ### “Goods” means all of the commodities, raw material, machinery and equipment, and/or other materials that the Supplier is required to supply to the Purchaser under the Contract. ### “Purchaser’s Country” is the country specified in the Special Conditions of Contract (SCC). ### “Purchaser” means the entity purchasing the Goods and Related Services, as specified in the SCC. ### “Related Services” means the services incidental to the supply of the goods, such as insurance, installation, training and initial maintenance and other such obligations of the Supplier under the Contract. ### “SCC” means the Special Conditions of Contract. ### “Subcontractor” means any person, private or government entity, or a combination of the above, to whom any part of the Goods to be supplied or execution of any part of the Related Services is subcontracted by the Supplier. ### “Supplier” means the person, private or government entity, or a combination of the above, whose Bid to perform the Contract has been accepted by the Purchaser and is named as such in the Contract Agreement. ### “The Project Site,” where applicable, means the place named in the SCC. |
| --- | --- |
| Contract Documents | - Subject to the order of precedence set forth in the Contract Agreement, all documents forming the Contract (and all parts thereof) are intended to be correlative, complementary, and mutually explanatory. The Contract Agreement shall be read as a whole. |
| Fraud and Corruption | 3.1 The Bank requires compliance with the Bank’s Anti-Corruption Guidelines and its prevailing sanctions policies and procedures as set forth in the WBG’s Sanctions Framework, as set forth in Appendix 1 to the GCC. 3.2 The Purchaser requires the Supplier to disclose any commissions or fees that may have been paid or are to be paid to agents or any other party with respect to the Bidding process or execution of the Contract. The information disclosed must include at least the name and address of the agent or other party, the amount and currency, and the purpose of the commission, gratuity or fee. |
| Interpretation | - If the context so requires it, singular means plural and vice versa. - Incoterms ### Unless inconsistent with any provision of the Contract, the meaning of any trade term and the rights and obligations of parties thereunder shall be as prescribed by Incoterms specified in the SCC. ### The terms EXW, CIP, FCA, CFR and other similar terms, when used, shall be governed by the rules prescribed in the current edition of Incoterms specified in the SCC and published by the International Chamber of Commerce in Paris, France. - Entire Agreement The Contract constitutes the entire agreement between the Purchaser and the Supplier and supersedes all communications, negotiations and agreements (whether written or oral) of the parties with respect thereto made prior to the date of Contract. - Amendment No amendment or other variation of the Contract shall be valid unless it is in writing, is dated, expressly refers to the Contract, and is signed by a duly authorized representative of each party thereto. - Nonwaiver ### Subject to GCC Sub-Clause 4.5(b) below, no relaxation, forbearance, delay, or indulgence by either party in enforcing any of the terms and conditions of the Contract or the granting of time by either party to the other shall prejudice, affect, or restrict the rights of that party under the Contract, neither shall any waiver by either party of any breach of Contract operate as waiver of any subsequent or continuing breach of Contract. ### Any waiver of a party’s rights, powers, or remedies under the Contract must be in writing, dated, and signed by an authorized representative of the party granting such waiver, and must specify the right and the extent to which it is being waived. - Severability If any provision or condition of the Contract is prohibited or rendered invalid or unenforceable, such prohibition, invalidity or unenforceability shall not affect the validity or enforceability of any other provisions and conditions of the Contract. |
| Language | - The Contract as well as all correspondence and documents relating to the Contract exchanged by the Supplier and the Purchaser, shall be written in the language specified in the SCC. Supporting documents and printed literature that are part of the Contract may be in another language provided they are accompanied by an accurate translation of the relevant passages in the language specified, in which case, for purposes of interpretation of the Contract, this translation shall govern. - The Supplier shall bear all costs of translation to the governing language and all risks of the accuracy of such translation, for documents provided by the Supplier. |
| Joint Venture, Consortium or Association | - If the Supplier is a joint venture, consortium, or association, all of the parties shall be jointly and severally liable to the Purchaser for the fulfillment of the provisions of the Contract and shall designate one party to act as a leader with authority to bind the joint venture, consortium, or association. The composition or the constitution of the joint venture, consortium, or association shall not be altered without the prior consent of the Purchaser. |
| Eligibility | - The Supplier and its Subcontractors shall have the nationality of an eligible country. A Supplier or Subcontractor shall be deemed to have the nationality of a country if it is a citizen or constituted, incorporated, or registered, and operates in conformity with the provisions of the laws of that country. - All Goods and Related Services to be supplied under the Contract and financed by the Bank shall have their origin in Eligible Countries. For the purpose of this Clause, origin means the country where the goods have been grown, mined, cultivated, produced, manufactured, or processed; or through manufacture, processing, or assembly, another commercially recognized article results that differs substantially in its basic characteristics from its components. |
| Notices | - Any notice given by one party to the other pursuant to the Contract shall be in writing to the address specified in the SCC. The term “in writing” means communicated in written form with proof of receipt. - A notice shall be effective when delivered or on the notice’s effective date, whichever is later. |
| Governing Law | - The Contract shall be governed by and interpreted in accordance with the laws of the Purchaser’s Country, unless otherwise specified in the SCC. - Throughout the execution of the Contract, the Supplier shall comply with the import of goods and services prohibitions in the Purchaser’s Country when: - as a matter of law or official regulations, the Borrower’s country prohibits commercial relations with that country; or - by an act of compliance with a decision of the United Nations Security Council taken under Chapter VII of the Charter of the United Nations, the Borrower’s Country prohibits any import of goods from that country or any payments to any country, person, or entity in that country. |
| Settlement of Disputes | - The Purchaser and the Supplier shall make every effort to resolve amicably by direct informal negotiation any disagreement or dispute arising between them under or in connection with the Contract. - If, after twenty-eight (28) days, the parties have failed to resolve their dispute or difference by such mutual consultation, then either the Purchaser or the Supplier may give notice to the other party of its intention to commence arbitration, as hereinafter provided, as to the matter in dispute, and no arbitration in respect of this matter may be commenced unless such notice is given. Any dispute or difference in respect of which a notice of intention to commence arbitration has been given in accordance with this Clause shall be finally settled by arbitration. Arbitration may be commenced prior to or after delivery of the Goods under the Contract. Arbitration proceedings shall be conducted in accordance with the rules of procedure specified in the SCC. - Notwithstanding any reference to arbitration herein, - the parties shall continue to perform their respective obligations under the Contract unless they otherwise agree; and - the Purchaser shall pay the Supplier any monies due the Supplier. |
| Inspections and Audit by the Bank | - The Supplier shall keep, and shall make all reasonable efforts to cause its Subcontractors and subconsultants to keep, accurate and systematic accounts and records in respect of the Goods in such form and details as will clearly identify relevant time changes and costs. - Pursuant to paragraph 2.2 e. of Appendix 1 to the General Conditions the Supplier shall permit and shall cause its agents (where declared or not), subcontractors, subconsultants, service providers, suppliers, and personnel, to permit, the Bank and/or persons appointed by the Bank to inspect the site and/or the accounts, records and other documents relating to the procurement process, selection and/or contract execution, and to have such accounts, records and other documents audited by auditors appointed by the Bank. The Supplier’s and its Subcontractors’ and subconsultants’ attention is drawn to Sub-Clause 3.1 (Fraud and Corruption) which provides, inter alia, that acts intended to materially impede the exercise of the Bank’s inspection and audit rights constitute a prohibited practice subject to contract termination (as well as to a determination of ineligibility pursuant to the Bank’s prevailing sanctions procedures). |
| Scope of Supply | 12.1 The Goods and Related Services to be supplied shall be as specified in the Schedule of Requirements. |
| Delivery and Documents | 13.1 Subject to GCC Sub-Clause 33.1, the Delivery of the Goods and Completion of the Related Services shall be in accordance with the Delivery and Completion Schedule specified in the Schedule of Requirements. The details of shipping and other documents to be furnished by the Supplier are specified in the SCC. |
| Supplier’s Responsibilities | - The Supplier shall supply all the Goods and Related Services included in the Scope of Supply in accordance with GCC Clause 12, and the Delivery and Completion Schedule, as per GCC Clause 13. - The Supplier, including its Subcontractors, shall not employ or engage forced labor or persons subject to trafficking, as described in GCC Sub-Clauses 14.3 and 14.4. - Forced labor consists of any work or service, not voluntarily performed, that is exacted from an individual under threat of force or penalty, and includes any kind of involuntary or compulsory labor, such as indentured labor, bonded labor or similar labor-contracting arrangements. - Trafficking in persons is defined as the recruitment, transportation, transfer, harbouring or receipt of persons by means of the threat or use of force or other forms of coercion, abduction, fraud, deception, abuse of power, or of a position of vulnerability, or of the giving or receiving of payments or benefits to achieve the consent of a person having control over another person, for the purposes of exploitation. - The Supplier, including its Subcontractors, shall not employ or engage a child under the age of 14 unless the national law specifies a higher age (the minimum age). - The Supplier, including its Subcontractors, shall not employ or engage a child between the minimum age and the age of 18 in a manner that is likely to be hazardous, or to interfere with, the child’s education, or to be harmful to the child’s health or physical, mental, spiritual, moral, or social development. - Work considered hazardous for children is work that, by its nature or the circumstances in which it is carried out, is likely to jeopardize the health, safety, or morals of children. Such work activities prohibited for children include work: - with exposure to physical, psychological or sexual abuse; - underground, underwater, working at heights or in confined spaces; - with dangerous machinery, equipment or tools, or involving handling or transport of heavy loads; - in unhealthy environments exposing children to hazardous substances, agents, or processes, or to temperatures, noise or vibration damaging to health; or - under difficult conditions such as work for long hours, during the night or in confinement on the premises of the employer. - The Supplier shall comply, and shall require its Subcontractors if any to comply, with all applicable health and safety regulations, laws, guidelines, and any other requirement stated in the Technical Specifications. - Pursuant to the SCC, the Supplier, including its Subcontractors/ suppliers/ manufacturers shall take all technical and organizational measures necessary to protect the information technology systems and data used in connection with the Contract. Without limiting the foregoing, the Supplier, including its Subcontractors/ suppliers/ manufacturers, shall use all reasonable efforts to establish, maintain, implement and comply with, reasonable information technology, information security, cyber security and data protection controls, policies and procedures, including oversight, access controls, encryption, technological and physical safeguards and business continuity/disaster recovery and security plans that are designed to protect against and prevent breach, destruction, loss, unauthorized distribution, use, access, disablement, misappropriation or modification, or other compromise or misuse of or relating to any information technology system or data used in connection with the Contract. - The Supplier shall comply with additional obligations as specified in the SCC. |
| Contract Price | 15.1 Prices charged by the Supplier for the Goods supplied and the Related Services performed under the Contract shall not vary from the prices quoted by the Supplier in its Bid, with the exception of any price adjustments authorized in the SCC. |
| Terms of Payment | 16.1 The Contract Price, including any Advance Payments, if applicable, shall be paid as specified in the SCC. 16.2 The Supplier’s request for payment shall be made to the Purchaser in writing, accompanied by invoices describing, as appropriate, the Goods delivered and Related Services performed, and by the documents submitted pursuant to GCC Clause 13 and upon fulfillment of all other obligations stipulated in the Contract. 16.3 Payments shall be made promptly by the Purchaser, but in no case later than sixty (60) days after submission of an invoice or request for payment by the Supplier, and after the Purchaser has accepted it. 16.4 The currencies in which payments shall be made to the Supplier under this Contract shall be those in which the Bid price is expressed. 16.5 In the event that the Purchaser fails to pay the Supplier any payment by its due date or within the period set forth in the SCC, the Purchaser shall pay to the Supplier interest on the amount of such delayed payment at the rate shown in the SCC, for the period of delay until payment has been made in full, whether before or after judgment or arbitrage award. |
| Taxes and Duties | 17.1 For goods manufactured outside the Purchaser’s Country, the Supplier shall be entirely responsible for all taxes, stamp duties, license fees, and other such levies imposed outside the Purchaser’s Country. 17.2 For goods Manufactured within the Purchaser’s Country, the Supplier shall be entirely responsible for all taxes, duties, license fees, etc., incurred until delivery of the contracted Goods to the Purchaser. 17.3 If any tax exemptions, reductions, allowances or privileges may be available to the Supplier in the Purchaser’s Country, the Purchaser shall use its best efforts to enable the Supplier to benefit from any such tax savings to the maximum allowable extent. |
| Performance Security | 18.1 If required as specified in the SCC, the Supplier shall, within twenty-eight (28) days of the notification of contract award, provide a performance security for the performance of the Contract in the amount specified in the SCC. 18.2 The proceeds of the Performance Security shall be payable to the Purchaser as compensation for any loss resulting from the Supplier’s failure to complete its obligations under the Contract. 18.3 As specified in the SCC, the Performance Security, if required, shall be denominated in the currency (ies) of the Contract, or in a freely convertible currency acceptable to the Purchaser; and shall be in one of the format stipulated by the Purchaser in the SCC, or in another format acceptable to the Purchaser. 18.4 The Performance Security shall be discharged by the Purchaser and returned to the Supplier not later than twenty-eight (28) days following the date of Completion of the Supplier’s performance obligations under the Contract, including any warranty obligations, unless specified otherwise in the SCC. |
| Copyright | 19.1 The copyright in all drawings, documents, and other materials containing data and information furnished to the Purchaser by the Supplier herein shall remain vested in the Supplier, or, if they are furnished to the Purchaser directly or through the Supplier by any third party, including suppliers of materials, the copyright in such materials shall remain vested in such third party. |
| Confidential Information | 20.1 The Purchaser and the Supplier shall keep confidential and shall not, without the written consent of the other party hereto, divulge to any third party any documents, data, or other information furnished directly or indirectly by the other party hereto in connection with the Contract, whether such information has been furnished prior to, during or following completion or termination of the Contract. Notwithstanding the above, the Supplier may furnish to its Subcontractor such documents, data, and other information it receives from the Purchaser to the extent required for the Subcontractor to perform its work under the Contract, in which event the Supplier shall obtain from such Subcontractor an undertaking of confidentiality similar to that imposed on the Supplier under GCC Clause 20. 20.2 The Purchaser shall not use such documents, data, and other information received from the Supplier for any purposes unrelated to the contract. Similarly, the Supplier shall not use such documents, data, and other information received from the Purchaser for any purpose other than the performance of the Contract. 20.3 The obligation of a party under GCC Sub-Clauses 20.1 and 20.2 above, however, shall not apply to information that: ### the Purchaser or Supplier need to share with the Bank or other institutions participating in the financing of the Contract; ### now or hereafter enters the public domain through no fault of that party; ### can be proven to have been possessed by that party at the time of disclosure and which was not previously obtained, directly or indirectly, from the other party; or ### otherwise lawfully becomes available to that party from a third party that has no obligation of confidentiality. 20.4 The above provisions of GCC Clause 20 shall not in any way modify any undertaking of confidentiality given by either of the parties hereto prior to the date of the Contract in respect of the Supply or any part thereof. 20.5 The provisions of GCC Clause 20 shall survive completion or termination, for whatever reason, of the Contract. |
| Subcontracting | 21.1 The Supplier shall notify the Purchaser in writing of all subcontracts awarded under the Contract if not already specified in the Bid. Notification by the Supplier, for addition of any Subcontractor not named in the Contract, shall also include the Subcontractor’s declaration in accordance with Appendix 2 to the GCC- Sexual exploitation and Abuse (SEA) and/or Sexual Harassment (SH) Performance Declaration. Such notification, in the original Bid or later shall not relieve the Supplier from any of its obligations, duties, responsibilities, or liability under the Contract. 21.2 Subcontracts shall comply with the provisions of GCC Clauses 3 and 7. |
| Specifications and Standards | 22.1 Technical Specifications and Drawings ### The Goods and Related Services supplied under this Contract shall conform to the technical specifications and standards mentioned in Section VI, Schedule of Requirements and, when no applicable standard is mentioned, the standard shall be equivalent or superior to the official standards whose application is appropriate to the Goods’ country of origin. ### The Supplier shall be entitled to disclaim responsibility for any design, data, drawing, specification or other document, or any modification thereof provided or designed by or on behalf of the Purchaser, by giving a notice of such disclaimer to the Purchaser. ### Wherever references are made in the Contract to codes and standards in accordance with which it shall be executed, the edition or the revised version of such codes and standards shall be those specified in the Schedule of Requirements. During Contract execution, any changes in any such codes and standards shall be applied only after approval by the Purchaser and shall be treated in accordance with GCC Clause 33. |
| Packing and Documents | 23.1 The Supplier shall provide such packing of the Goods as is required to prevent their damage or deterioration during transit to their final destination, as indicated in the Contract. During transit, the packing shall be sufficient to withstand, without limitation, rough handling and exposure to extreme temperatures, salt and precipitation, and open storage. Packing case size and weights shall take into consideration, where appropriate, the remoteness of the goods’ final destination and the absence of heavy handling facilities at all points in transit. 23.2 The packing, marking, and documentation within and outside the packages shall comply strictly with such special requirements as shall be expressly provided for in the Contract, including additional requirements, if any, specified in the SCC, and in any other instructions ordered by the Purchaser. |
| Insurance | 24.1 Unless otherwise specified in the SCC, the Goods supplied under the Contract shall be fully insured—in a freely convertible currency from an eligible country—against loss or damage incidental to manufacture or acquisition, transportation, storage, and delivery, in accordance with the applicable Incoterms or in the manner specified in the SCC. |
| Transportation and Incidental Services | 25.1 Unless otherwise specified in the SCC, responsibility for arranging transportation of the Goods shall be in accordance with the specified Incoterms. |
|  | 25.2 The Supplier may be required to provide any or all of the following services, including additional services, if any, specified in SCC: (a) performance or supervision of on-site assembly and/or start-up of the supplied Goods; (b) furnishing of tools required for assembly and/or maintenance of the supplied Goods; (c) furnishing of a detailed operations and maintenance manual for each appropriate unit of the supplied Goods; (d) performance or supervision or maintenance and/or repair of the supplied Goods, for a period of time agreed by the parties, provided that this service shall not relieve the Supplier of any warranty obligations under this Contract; and (e) training of the Purchaser’s personnel, at the Supplier’s plant and/or on-site, in assembly, start-up, operation, maintenance, and/or repair of the supplied Goods. 25.3 Prices charged by the Supplier for incidental services, if not included in the Contract Price for the Goods, shall be agreed upon in advance by the parties and shall not exceed the prevailing rates charged to other parties by the Supplier for similar services |
| Inspections and Tests | 26.1 The Supplier shall at its own expense and at no cost to the Purchaser carry out all such tests and/or inspections of the Goods and Related Services as are specified in the SCC. 26.2 The inspections and tests may be conducted on the premises of the Supplier or its Subcontractor, at point of delivery, and/or at the Goods’ final destination, or in another place in the Purchaser’s Country as specified in the SCC. Subject to GCC Sub-Clause 26.3, if conducted on the premises of the Supplier or its Subcontractor, all reasonable facilities and assistance, including access to drawings and production data, shall be furnished to the inspectors at no charge to the Purchaser. 26.3 The Purchaser or its designated representative shall be entitled to attend the tests and/or inspections referred to in GCC Sub-Clause 26.2, provided that the Purchaser bear all of its own costs and expenses incurred in connection with such attendance including, but not limited to, all traveling and board and lodging expenses. 26.4 Whenever the Supplier is ready to carry out any such test and inspection, it shall give a reasonable advance notice, including the place and time, to the Purchaser. The Supplier shall obtain from any relevant third party or manufacturer any necessary permission or consent to enable the Purchaser or its designated representative to attend the test and/or inspection. 26.5 The Purchaser may require the Supplier to carry out any test and/or inspection not required by the Contract but deemed necessary to verify that the characteristics and performance of the Goods comply with the technical specifications codes and standards under the Contract, provided that the Supplier’s reasonable costs and expenses incurred in the carrying out of such test and/or inspection shall be added to the Contract Price. Further, if such test and/or inspection impedes the progress of manufacturing and/or the Supplier’s performance of its other obligations under the Contract, due allowance will be made in respect of the Delivery Dates and Completion Dates and the other obligations so affected. 26.6 The Supplier shall provide the Purchaser with a report of the results of any such test and/or inspection. 26.7 The Purchaser may reject any Goods or any part thereof that fail to pass any test and/or inspection or do not conform to the specifications. The Supplier shall either rectify or replace such rejected Goods or parts thereof or make alterations necessary to meet the specifications at no cost to the Purchaser, and shall repeat the test and/or inspection, at no cost to the Purchaser, upon giving a notice pursuant to GCC Sub-Clause 26.4. 26.8 The Supplier agrees that neither the execution of a test and/or inspection of the Goods or any part thereof, nor the attendance by the Purchaser or its representative, nor the issue of any report pursuant to GCC Sub-Clause 26.6, shall release the Supplier from any warranties or other obligations under the Contract. |
| Liquidated Damages | 27.1 Except as provided under GCC Clause 32, if the Supplier fails to deliver any or all of the Goods by the Date(s) of delivery or perform the Related Services within the period specified in the Contract, the Purchaser may without prejudice to all its other remedies under the Contract, deduct from the Contract Price, as liquidated damages, a sum equivalent to the percentage specified in the SCC of the delivered price of the delayed Goods or unperformed Services for each week or part thereof of delay until actual delivery or performance, up to a maximum deduction of the percentage specified in those SCC. Once the maximum is reached, the Purchaser may terminate the Contract pursuant to GCC Clause 35. |
| Warranty | 28.1 The Supplier warrants that all the Goods are new, unused, and of the most recent or current models, and that they incorporate all recent improvements in design and materials, unless provided otherwise in the Contract. 28.2 Subject to GCC Sub-Clause 22.1(b), the Supplier further warrants that the Goods shall be free from defects arising from any act or omission of the Supplier or arising from design, materials, and workmanship, under normal use in the conditions prevailing in the country of final destination. 28.3 Unless otherwise specified in the SCC, the warranty shall remain valid for twelve (12) months after the Goods, or any portion thereof as the case may be, have been delivered to and accepted at the final destination indicated in the SCC, or for eighteen (18) months after the date of shipment from the port or place of loading in the country of origin, whichever period concludes earlier. 28.4 The Purchaser shall give notice to the Supplier stating the nature of any such defects together with all available evidence thereof, promptly following the discovery thereof. The Purchaser shall afford all reasonable opportunity for the Supplier to inspect such defects. 28.5 Upon receipt of such notice, the Supplier shall, within the period specified in the SCC, expeditiously repair or replace the defective Goods or parts thereof, at no cost to the Purchaser. 28.6 If having been notified, the Supplier fails to remedy the defect within the period specified in the SCC, the Purchaser may proceed to take within a reasonable period such remedial action as may be necessary, at the Supplier’s risk and expense and without prejudice to any other rights which the Purchaser may have against the Supplier under the Contract. |
| Patent Indemnity | 29.1 The Supplier shall, subject to the Purchaser’s compliance with GCC Sub-Clause 29.2, indemnify and hold harmless the Purchaser and its employees and officers from and against any and all suits, actions or administrative proceedings, claims, demands, losses, damages, costs, and expenses of any nature, including attorney’s fees and expenses, which the Purchaser may suffer as a result of any infringement or alleged infringement of any patent, utility model, registered design, trademark, copyright, or other intellectual property right registered or otherwise existing at the date of the Contract by reason of: ### the installation of the Goods by the Supplier or the use of the Goods in the country where the Site is located; and ### the sale in any country of the products produced by the Goods. ### Such indemnity shall not cover any use of the Goods or any part thereof other than for the purpose indicated by or to be reasonably inferred from the Contract, neither any infringement resulting from the use of the Goods or any part thereof, or any products produced thereby in association or combination with any other equipment, plant, or materials not supplied by the Supplier, pursuant to the Contract. 29.2 If any proceedings are brought or any claim is made against the Purchaser arising out of the matters referred to in GCC Sub-Clause 29.1, the Purchaser shall promptly give the Supplier a notice thereof, and the Supplier may at its own expense and in the Purchaser’s name conduct such proceedings or claim and any negotiations for the settlement of any such proceedings or claim. 29.3 If the Supplier fails to notify the Purchaser within twenty-eight (28) days after receipt of such notice that it intends to conduct any such proceedings or claim, then the Purchaser shall be free to conduct the same on its own behalf. 29.4 The Purchaser shall, at the Supplier’s request, afford all available assistance to the Supplier in conducting such proceedings or claim, and shall be reimbursed by the Supplier for all reasonable expenses incurred in so doing. 29.5 The Purchaser shall indemnify and hold harmless the Supplier and its employees, officers, and Subcontractors from and against any and all suits, actions or administrative proceedings, claims, demands, losses, damages, costs, and expenses of any nature, including attorney’s fees and expenses, which the Supplier may suffer as a result of any infringement or alleged infringement of any patent, utility model, registered design, trademark, copyright, or other intellectual property right registered or otherwise existing at the date of the Contract arising out of or in connection with any design, data, drawing, specification, or other documents or materials provided or designed by or on behalf of the Purchaser. |
| Limitation of Liability | 30.1 Except in cases of criminal negligence or willful misconduct, (a) the Supplier shall not be liable to the Purchaser, whether in contract, tort, or otherwise, for any indirect or consequential loss or damage, loss of use, loss of production, or loss of profits or interest costs, provided that this exclusion shall not apply to any obligation of the Supplier to pay liquidated damages to the Purchaser and (b) the aggregate liability of the Supplier to the Purchaser, whether under the Contract, in tort or otherwise, shall not exceed the total Contract Price, provided that this limitation shall not apply to the cost of repairing or replacing defective equipment, or to any obligation of the supplier to indemnify the Purchaser with respect to patent infringement |
| Change in Laws and Regulations | 31.1 Unless otherwise specified in the Contract, if after the date of 28 days prior to date of Bid submission, any law, regulation, ordinance, order or bylaw having the force of law is enacted, promulgated, abrogated, or changed in the place of the Purchaser’s Country where the Site is located (which shall be deemed to include any change in interpretation or application by the competent authorities) that subsequently affects the Delivery Date and/or the Contract Price, then such Delivery Date and/or Contract Price shall be correspondingly increased or decreased, to the extent that the Supplier has thereby been affected in the performance of any of its obligations under the Contract. Notwithstanding the foregoing, such additional or reduced cost shall not be separately paid or credited if the same has already been accounted for in the price adjustment provisions where applicable, in accordance with GCC Clause 15. |
| Force Majeure | 32.1 The Supplier shall not be liable for forfeiture of its Performance Security, liquidated damages, or termination for default if and to the extent that it’s delay in performance or other failure to perform its obligations under the Contract is the result of an event of Force Majeure. 32.2 For purposes of this Clause, “Force Majeure” means an event or situation beyond the control of the Supplier that is not foreseeable, is unavoidable, and its origin is not due to negligence or lack of care on the part of the Supplier. Such events may include, but not be limited to, acts of the Purchaser in its sovereign capacity, wars or revolutions, fires, floods, epidemics, quarantine restrictions, and freight embargoes. 32.3 If a Force Majeure situation arises, the Supplier shall promptly notify the Purchaser in writing of such condition and the cause thereof. Unless otherwise directed by the Purchaser in writing, the Supplier shall continue to perform its obligations under the Contract as far as is reasonably practical, and shall seek all reasonable alternative means for performance not prevented by the Force Majeure event. |
| Change Orders and Contract Amendments | - The Purchaser may at any time order the Supplier through notice in accordance GCC Clause 8, to make changes within the general scope of the Contract in any one or more of the following: ### drawings, designs, or specifications, where Goods to be furnished under the Contract are to be specifically manufactured for the Purchaser; ### the method of shipment or packing; ### the place of delivery; and ### the Related Services to be provided by the Supplier. - If any such change causes an increase or decrease in the cost of, or the time required for, the Supplier’s performance of any provisions under the Contract, an equitable adjustment shall be made in the Contract Price or in the Delivery/Completion Schedule, or both, and the Contract shall accordingly be amended. Any claims by the Supplier for adjustment under this Clause must be asserted within twenty-eight (28) days from the date of the Supplier’s receipt of the Purchaser’s change order. - Prices to be charged by the Supplier for any Related Services that might be needed but which were not included in the Contract shall be agreed upon in advance by the parties and shall not exceed the prevailing rates charged to other parties by the Supplier for similar services. - Value Engineering: The Supplier may prepare, at its own cost, a value engineering proposal at any time during the performance of the contract. The value engineering proposal shall, at a minimum, include the following; - the proposed change(s), and a description of the difference to the existing contract requirements; - a full cost/benefit analysis of the proposed change(s) including a description and estimate of costs (including life cycle costs) the Purchaser may incur in implementing the value engineering proposal; and - a description of any effect(s) of the change on performance/functionality. The Purchaser may accept the value engineering proposal if the proposal demonstrates benefits that: - accelerates the delivery period; or - reduces the Contract Price or the life cycle costs to the Purchaser; or - improves the quality, efficiency or sustainability of the Goods; or - yields any other benefits to the Purchaser, without compromising the necessary functions of the Facilities. If the value engineering proposal is approved by the Purchaser and results in: - a reduction of the Contract Price; the amount to be paid to the Supplier shall be the percentage specified in the PCC of the reduction in the Contract Price; or - an increase in the Contract Price; but results in a reduction in life cycle costs due to any benefit described in (a) to (d) above, the amount to be paid to the Supplier shall be the full increase in the Contract Price. - Subject to the above, no variation in or modification of the terms of the Contract shall be made except by written amendment signed by the parties. |
| Extensions of Time | 34.1 If at any time during performance of the Contract, the Supplier or its subcontractors should encounter conditions impeding timely delivery of the Goods or completion of Related Services pursuant to GCC Clause 13, the Supplier shall promptly notify the Purchaser in writing of the delay, its likely duration, and its cause. As soon as practicable after receipt of the Supplier’s notice, the Purchaser shall evaluate the situation and may at its discretion extend the Supplier’s time for performance, in which case the extension shall be ratified by the parties by amendment of the Contract. 34.2 Except in case of Force Majeure, as provided under GCC Clause 32, a delay by the Supplier in the performance of its Delivery and Completion obligations shall render the Supplier liable to the imposition of liquidated damages pursuant to GCC Clause 26, unless an extension of time is agreed upon, pursuant to GCC Sub-Clause 34.1. |
| Termination | 35.1 Termination for Default ### The Purchaser, without prejudice to any other remedy for breach of Contract, by written notice of default sent to the Supplier, may terminate the Contract in whole or in part: #### if the Supplier fails to deliver any or all of the Goods within the period specified in the Contract, or within any extension thereof granted by the Purchaser pursuant to GCC Clause 34; #### if the Supplier fails to perform any other obligation under the Contract; or #### if the Supplier, in the judgment of the Purchaser has engaged in Fraud and Corruption, as defined in paragraph 2.2 a of the Appendix 1 to the GCC, in competing for or in executing the Contract. ### In the event the Purchaser terminates the Contract in whole or in part, pursuant to GCC Clause 35.1(a), the Purchaser may procure, upon such terms and in such manner as it deems appropriate, Goods or Related Services similar to those undelivered or not performed, and the Supplier shall be liable to the Purchaser for any additional costs for such similar Goods or Related Services. However, the Supplier shall continue performance of the Contract to the extent not terminated. 35.2 Termination for Insolvency. ### The Purchaser may at any time terminate the Contract by giving notice to the Supplier if the Supplier becomes bankrupt or otherwise insolvent. In such event, termination will be without compensation to the Supplier, provided that such termination will not prejudice or affect any right of action or remedy that has accrued or will accrue thereafter to the Purchaser 35.3 Termination for Convenience. ### The Purchaser, by notice sent to the Supplier, may terminate the Contract, in whole or in part, at any time for its convenience. The notice of termination shall specify that termination is for the Purchaser’s convenience, the extent to which performance of the Supplier under the Contract is terminated, and the date upon which such termination becomes effective. ### The Goods that are complete and ready for shipment within twenty-eight (28) days after the Supplier’s receipt of notice of termination shall be accepted by the Purchaser at the Contract terms and prices. For the remaining Goods, the Purchaser may elect: #### to have any portion completed and delivered at the Contract terms and prices; and/or #### to cancel the remainder and pay to the Supplier an agreed amount for partially completed Goods and Related Services and for materials and parts previously procured by the Supplier. |
| Assignment | 36.1 Neither the Purchaser nor the Supplier shall assign, in whole or in part, their obligations under this Contract, except with prior written consent of the other party. |
| Export Restriction | 37.1 Notwithstanding any obligation under the Contract to complete all export formalities, any export restrictions attributable to the Purchaser, to the country of the Purchaser, or to the use of the products/goods, systems or services to be supplied, which arise from trade regulations from a country supplying those products/goods, systems or services, and which substantially impede the Supplier from meeting its obligations under the Contract, shall release the Supplier from the obligation to provide deliveries or services, always provided, however, that the Supplier can demonstrate to the satisfaction of the Purchaser and of the Bank that it has completed all formalities in a timely manner, including applying for permits, authorizations and licenses necessary for the export of the products/goods, systems or services under the terms of the Contract. Termination of the Contract on this basis shall be for the Purchaser’s convenience pursuant to Sub-Clause 35.3. |

APPENDIX 1

(Text in this Appendix shall not be modified)

Fraud and Corruption

- Purpose

  - The Bank’s Anti-Corruption Guidelines and this annex apply with respect to procurement under Bank Investment Project Financing operations.

- Requirements

- The Bank requires that Borrowers (including beneficiaries of Bank financing); bidders (applicants/proposers), consultants, contractors and suppliers; any sub-contractors, sub-consultants, service providers or suppliers; any agents (whether declared or not); and any of their personnel, observe the highest standard of ethics during the procurement process, selection and contract execution of Bank-financed contracts, and refrain from Fraud and Corruption.

- To this end, the Bank:

- Defines, for the purposes of this provision, the terms set forth below as follows:

- “corrupt practice” is the offering, giving, receiving, or soliciting, directly or indirectly, of anything of value to influence improperly the actions of another party;

- “fraudulent practice” is any act or omission, including misrepresentation, that knowingly or recklessly misleads, or attempts to mislead, a party to obtain financial or other benefit or to avoid an obligation;

- “collusive practice” is an arrangement between two or more parties designed to achieve an improper purpose, including to influence improperly the actions of another party;

- “coercive practice” is impairing or harming, or threatening to impair or harm, directly or indirectly, any party or the property of the party to influence improperly the actions of a party;

- “obstructive practice” is:

- deliberately destroying, falsifying, altering, or concealing of evidence material to the investigation or making false statements to investigators in order to materially impede a Bank investigation into allegations of a corrupt, fraudulent, coercive, or collusive practice; and/or threatening, harassing, or intimidating any party to prevent it from disclosing its knowledge of matters relevant to the investigation or from pursuing the investigation; or

- acts intended to materially impede the exercise of the Bank’s inspection and audit rights provided for under paragraph 2.2 e. below.

- Rejects a proposal for award if the Bank determines that the firm or individual recommended for award, any of its personnel, or its agents, or its sub-consultants, sub-contractors, service providers, suppliers and/ or their employees, has, directly or indirectly, engaged in corrupt, fraudulent, collusive, coercive, or obstructive practices in competing for the contract in question;

- In addition to the legal remedies set out in the relevant Legal Agreement, may take other appropriate actions, including declaring misprocurement, if the Bank determines at any time that representatives of the Borrower or of a recipient of any part of the proceeds of the loan engaged in corrupt, fraudulent, collusive, coercive, or obstructive practices during the procurement process, selection and/or execution of the contract in question,  without the Borrower having taken timely and appropriate action satisfactory to the Bank to address such practices when they occur, including by failing to inform the Bank in a timely manner at the time  they knew of the practices;

- Pursuant to the Bank’s Anti-Corruption Guidelines, and in accordance with the Bank’s prevailing sanctions policies and procedures, may sanction a firm or individual, either indefinitely or for a stated period of time, including by publicly declaring such firm or individual ineligible (i) to be awarded or otherwise benefit from a Bank-financed contract, financially or in any other manner; (ii) to be a nominated sub-contractor, consultant, manufacturer or supplier, or service provider of an otherwise eligible firm being awarded a Bank-financed contract; and (iii) to receive the proceeds of any loan made by the Bank or otherwise to participate further in the preparation or implementation of any Bank-financed project;

- Requires that a clause be included in bidding/request for proposals documents and in contracts financed by a Bank loan, requiring (i) bidders (applicants/proposers), consultants, contractors, and suppliers, and their sub-contractors, sub-consultants, service providers, suppliers, agents, personnel, permit the Bank to inspect all accounts, records and other documents relating to the procurement process, selection and/or contract execution, and to have them audited by auditors appointed by the Bank.

APPENDIX 2

Sexual Exploitation and Abuse (SEA) and/or Sexual Harassment (SH) Performance Declaration for Subcontractors*

[The following table shall be filled in by each subcontractor proposed by the Supplier, that was not named in the Contract]

Subcontractor’s Name: [insert full name]

Date: [insert day, month, year]
Contract reference [insert contract reference]
Page [insert page number] of [insert total number] pages

SEA and/or SH Declaration

We: (a) have not been subject to disqualification by the Bank for non-compliance with SEA/ SH obligations. (b) are subject to disqualification by the Bank for non-compliance with SEA/ SH obligations. (c) had been subject to disqualification by the Bank for non-compliance with SEA/ SH obligations, and were removed from the disqualification list. An arbitral award on the disqualification case has been made in our favor.

[If (c) above is applicable, attach evidence of an arbitral award reversing the findings on the issues underlying the disqualification.]

Period of disqualification: From: _______________ To: ________________

Name of the Subcontractor

Name of the person duly authorized to sign on behalf of the Subcontractor	_______

Title of the person signing on behalf of the Subcontractor	______________________

Signature of the person named above	______________________

Date signed ________________________________ day of ___________________, _____

Countersignature of authorized representative of the Supplier:

Signature: ________________________________________________________

Date signed ________________________________ day of ___________________, _____

| Section IX - Special Conditions of Contract |  |
| --- | --- |
| The following Special Conditions of Contract (SCC) shall supplement and / or amend the General Conditions of Contract (GCC). Whenever there is a conflict, the provisions herein shall prevail over those in the GCC. |  |
| GCC 1.1(i) | The Purchaser’s Country is: Federal Republic of Somalia |
| GCC 1.1(j) | The Purchaser is: Ministry of Finance of the Federal Government of Somalia. |
| GCC 1.1 (o) | The Project Site(s)/Final Destination(s) is/are: Somali Bureau of Standards (SoBS) – National Reference Laboratory (NRL), Mogadishu, Federal Republic of Somalia. The exact delivery address and site details shall be confirmed at contract signature. |
| GCC 1.1 (p) | ### The term SEA/SH where used in the Contract has the following meaning: ### “Sexual Exploitation and Abuse” “(SEA)” means the following: Sexual Exploitation is defined as any actual or attempted abuse of position of vulnerability, differential power or trust, for sexual purposes, including, but not limited to, profiting monetarily, socially or politically from the sexual exploitation of another. Sexual Abuse is defined as the actual or threatened physical intrusion of a sexual nature, whether by force or under unequal or coercive conditions. ### “Sexual Harassment” “(SH)” is defined as unwelcome sexual advances, requests for sexual favors, and other verbal or physical conduct of a sexual nature by supplier’s personnel with other supplier’s, or purchaser’s personnel. |
| GCC 1.1 (additional definition) | “Operational Acceptance” means the successful completion of delivery, installation, commissioning, calibration/verification and IQ/OQ/PQ qualification (where applicable), safety checks, cybersecurity configuration and backup/restore testing for networked components, and completion of training and handover documentation, including the ESHS close-out records (waste and packaging removal evidence, Safety Data Sheet (SDS) register handover and incident log), for the relevant Goods and Related Services, as evidenced by the Operational Acceptance Certificate(s) issued by the Purchaser. |
| GCC 4.2 (a) | The meaning of the trade terms shall be as prescribed by Incoterms. If the meaning of any trade term and the rights and obligations of the parties thereunder shall not be as prescribed by Incoterms, they shall be as prescribed by: Incoterms shall govern trade terms, rights and obligations |
| GCC 4.2 (b) | The version edition of Incoterms shall be Incoterms 2020 |
| GCC 5.1 | The language shall be: English language |
| GCC 8.1 | For notices, the Purchaser’s address shall be: Attention: Procurement Specialist Address: Ministry of Finance building DRIVE Project. Floor/ Room number: 3rd floor Email: bid.drive@piu.mof.gov.so, and copy procurement.drive@piu.mof.gov.so, City: Mogadishu Country: Somalia |
| GCC 9.1 | The governing law shall be the law of: Somalia |
| GCC 10.2 | The rules of procedure for arbitration proceedings pursuant to GCC Clause 10.2 shall be as follows: (a) Contracts with foreign supplier; GCC 10.2 (a)—Any dispute, controversy or claim arising out of or relating to this Contract, or breach, termination or invalidity thereof, shall be settled by arbitration in accordance with the UNCITRAL Arbitration Rules as at present in force. The appointing authority shall be the Secretary-General of the Permanent Court of Arbitration. The seat/place of arbitration shall be Nairobi, Kenya. The language of arbitration shall be English. The number of arbitrators shall be three (3), except that disputes with an amount in controversy below USD 500,000 may be determined by one (1) arbitrator. (b) Contracts with Supplier national of the Purchaser’s Country: In the case of a dispute between the Purchaser and a Supplier who is a national of the Purchaser’s Country, the dispute shall be referred to adjudication or arbitration in accordance with the laws of the Purchaser’s Country. |
| GCC 13.1 | Details of Shipping and other Documents to be furnished by the Supplier are; signed commercial invoice; packing list; a negotiable bill of lading, an airway bill (where applicable) or a road consignment note; insurance certificate; certificate of origin; manufacturer’s certificate of conformity; manufacturer’s or Supplier’s warranty certificate; inspection/test certificate issued by the nominated inspection agency (where applicable); calibration/verification certificates (where applicable); software license and activation documentation (where applicable); Safety Data Sheets and certificates of analysis for chemicals, reagents, gases and reference materials; cold-chain temperature records for temperature-controlled items; import/customs documentation; delivery note; and any other documents required under Section VII. The above documents shall be received by the Purchaser before arrival of the Goods and, if not received, the Supplier will be responsible for any consequent expenses. |
| GCC 14.9 | Cyber Security: applies to LIMS/QMS software, networked instruments, computers, servers/NAS, data storage, remote-support functions, user-access systems and other IT-enabled components supplied under the Contract. Cyber security does not apply to purely standalone, non-networked equipment except to the extent software, firmware, removable media or data transfer is used. |
| GCC 14.10 | Additional ESHS/OHS, cybersecurity and supply-chain obligations: (a) The Supplier shall comply with the World Bank Environmental and Social Standards applicable to the Project and with the DRIVE Project E&S instruments — the Environmental and Social Commitment Plan (ESCP), the Environmental and Social Management Framework (ESMF), the Stakeholder Engagement Plan (SEP) and the E&S procedures of the Project Implementation Manual — as updated or supplemented by site-specific requirements issued by the Purchaser. Before mobilization to site, the Supplier shall submit and implement an ESHS/OHS management plan acceptable to the Purchaser, including occupational health and safety controls, PPE, safe delivery/offloading/installation method statements, hazardous-material management, spill response, waste/e-waste management, a worker Code of Conduct with SEA/SH prevention, and grievance and incident reporting arrangements. (b) The Supplier shall notify the Purchaser immediately, and in any event within twenty-four (24) hours after becoming aware, of any serious incident related to the Contract (including injury, fire, spill, exposure, security incident, SEA/SH allegation, labour incident, environmental release, cold-chain failure or cybersecurity incident), followed by a written incident report within forty-eight (48) hours. (c) The Supplier shall implement the cybersecurity controls specified in Section VII for all software, LIMS, networked equipment and data-transfer interfaces, including access controls, administrator-credential handover, audit logging, patching, backup/restore testing, vulnerability remediation and breach notification. (d) The Supplier shall manage supply-chain risks for critical and long-lead equipment, spare parts, reagents/standards and software licenses, shall promptly notify the Purchaser of risks affecting delivery, warranty, service support, quality or contract completion, and shall remain fully responsible for the performance of its OEMs and subcontractors for the whole Contract. |
| GCC 15.1 | The prices charged for the Goods supplied and the related Services performed shall not, be adjustable. |
| GCC 16.1 | GCC 16.1—The method and conditions of payment to be made to the Supplier under this Contract shall be as follows: Payment for Goods supplied from abroad: Payment of foreign currency portion shall be made in United States Dollars (USD) in the following manner: (i) Advance Payment: Ten (10) percent of the Contract Price shall be paid within forty-five (45) days of signing of the Contract, and upon submission of claim and a bank guarantee for equivalent amount valid until the Goods are delivered and, in the form, provided in the bidding document or another form acceptable to the Purchaser. (ii) Seventy (70) percent of the Contract Price for the Goods shall be paid within forty-five (45) days of delivery of all Goods DDP (Incoterms 2020) at the Final Destination, upon submission of the shipping documents and the Purchaser’s certificate of delivery. The advance payment under (i) shall be recovered by deduction of its full amount from this payment. (iii) Twenty (20) percent of the Contract Price shall be paid within forty-five (45) days of successful completion of installation, commissioning, calibration/verification and IQ/OQ/PQ qualification, training and handover documentation, and issuance by the Purchaser of the Operational Acceptance Certificate(s) for all instruments and Related Services. Payment of local currency portion shall be made in United States Dollars within forty-five (45) days of presentation of claim supported by a certificate from the Purchaser declaring that the Goods have been delivered and that all other contracted Services have been performed. |
| GCC 16.5 | The payment-delay period after which the Purchaser shall pay interest to the supplier shall be Sixty (60) days. The interest rate that shall be applied is 2% |
| GCC 18.1 | A Performance Security shall be required The Performance Security shall be denominated in US Dollars for an amount equal to ten (10) percent of the Contract Price. During the Warranty Period (i.e., after Operational Acceptance of the System), the Performance Security shall be reduced to Five (5) percent of the Contract Price. |
| GCC 18.3 | Performance Security shall be in the form of: a Demand Guarantee If required, the Performance security shall be denominated in the currencies of payment of the Contract, in accordance with their portions of the Contract Price |
| GCC 18.4 | Discharge of the Performance Security shall take place: not later than twenty-eight (28) days after issuance of the Operational Acceptance Certificate for all Goods and Related Services; such discharge shall not release the Supplier from its warranty obligations or any other obligations surviving under the Contract. |
| GCC 23.2 | The packing, marking and documentation within and outside the packages shall be: RFB No. SO-MOF-561587-GO-RFB Contract Title: Supply, Delivery, Installation and Commissioning of Laboratory Equipment for the National Reference Laboratory (NRL) of the Somali Bureau of Standards (SoBS) Purchaser: Ministry of Finance, Federal Republic of Somalia Project: DRIVE Project (P176517) Project Site: SoBS – NRL, Mogadishu, Federal Republic of Somalia package number; item description; gross/net weight; handling instructions (fragile / keep dry / this side up, as applicable); storage temperature and humidity requirements where applicable; hazardous-material labels where applicable; and serial/model numbers where applicable. Packing shall be suitable for international transport, inland transport, handling and temporary storage in Somalia. |
| GCC 24.1 | The Supplier shall provide insurance covering, as a minimum: transit and inland transportation; marine/air/cargo risks; unloading and internal movement; installation all-risks; cold-chain and temperature-controlled items; equipment damage during installation and commissioning; third-party liability with coverage of not less than USD 500,000 per occurrence; and product liability for the supplied equipment. The insurance shall remain in force until Operational Acceptance of the relevant Goods and completion of the Related Services. |
| GCC 25.1 | Responsibility for transportation of the Goods shall be as specified in the Incoterms. |
| GCC 25.2 | Incidental services to be provided are: the Related Services specified in Section VII, Schedule of Requirements (List of Related Services and Completion Schedule), including pre-delivery site-readiness assessment, installation, commissioning and IQ/OQ/PQ qualification, training, proficiency testing enrolment, first-year external calibration services, handover documentation package, and cybersecurity hardening and LIMS/interface handover. |
| GCC 26.1 | The inspections and tests shall be: as specified in Section VII, Schedule of Requirements (Inspections and Tests), comprising pre-shipment verification and post-delivery inspection, testing and acceptance, including IQ/OQ/PQ qualification of each instrument. Factory Acceptance Tests (FAT) shall be conducted where required by the Purchaser, and Site Acceptance Tests (SAT) at the Final Destination. The Supplier shall provide test protocols for the Purchaser’s review and approval before testing and shall bear the cost of tests required under the Contract, including retesting required because of Supplier-attributable failure. |
| GCC 26.2 | The Inspections and tests shall be conducted at: - the manufacturer’s premises, where Factory Acceptance Tests (FAT) are required; and - the Final Destination: Somali Bureau of Standards (SoBS) – National Reference Laboratory (NRL), Mogadishu, Federal Republic of Somalia. |
| GCC 27.1 | The liquidated damages shall be: 0.5% of the delivered price of the delayed Goods or delayed Related Services for each week or part thereof of delay. Where the delayed item prevents operational use of an integrated subsystem or of the Contract package as a whole, liquidated damages may be applied to the affected subsystem or the affected portion of the Contract Price identified in the Price Schedules. |
| GCC 27.1 | The maximum amount of liquidated damages shall be: 10% |
| GCC 28.3 | The period of validity of the Warranty shall be: twenty-four (24) months from the date of Operational Acceptance of each instrument, or thirty (30) months from the date of delivery, whichever occurs first, unless a longer manufacturer warranty is offered in the Bid For purposes of the Warranty, the place(s) of final destination(s) shall be: Somali Bureau of Standards (SoBS) – National Reference Laboratory (NRL), Mogadishu, Federal Republic of Somalia GCC 28.3—In partial modification of the provisions, the warranty period shall be twenty-four (24) months from the date of Operational Acceptance of each instrument, or thirty (30) months from the date of delivery, whichever occurs first, unless a longer manufacturer warranty is offered in the Bid. The warranty period for any repaired or replaced item shall be extended by the period during which the item was unavailable for use due to a Supplier-attributable defect. During the warranty period the Supplier shall: provide parts and labour at no cost to the Purchaser; acknowledge and commence remote troubleshooting within two (2) Business Days of notification of a defect; attend on-site within five (5) Business Days for critical equipment and within ten (10) Business Days for non-critical equipment, where local or regional service support is required and site access and security conditions permit; restore the equipment or agree a corrective action plan with the Purchaser within thirty (30) calendar days; carry out at least two (2) preventive maintenance visits per year; maintain a service log and submit warranty and maintenance reports to the Purchaser; and ensure the availability of spare parts and consumables for at least five (5) years from Operational Acceptance. Downtime exceeding thirty (30) consecutive days for critical equipment shall trigger an escalation plan and, where feasible, the provision of loan or replacement equipment or other mitigation acceptable to the Purchaser. The Supplier shall, in addition, comply with the performance and/or consumption guarantees specified under the Contract. If, for reasons attributable to the Supplier, these guarantees are not attained in whole or in part, the Supplier shall, at its discretion, either: (a) make such changes, modifications, and/or additions to the Goods or any part thereof as may be necessary in order to attain the contractual guarantees specified in the Contract at its own cost and expense and to carry out further performance tests in accordance with GCC 26.7, or (b) pay liquidated damages to the Purchaser with respect to the failure to meet the contractual guarantees. The rate of these liquidated damages shall be 10%. |
| GCC 28.5 & 28.6 | The period for repair or replacement shall be: 14 days. |
| GCC 33.4 | If the value engineering proposal is approved by the Purchaser the amount to be paid to the Supplier shall be 30% |

Attachment: Price Adjustment Formula:  NA

If in accordance with GCC 15.1, prices shall be adjustable, the following method shall be used to calculate the price adjustment:

15.1 	Prices payable to the Supplier, as stated in the Contract, shall be subject to adjustment during performance of the Contract to reflect changes in the cost of labor and material components in accordance with the formula:

P1 = P0 [a + bL1 + cM1] - P0

L0	 M0

a+b+c = 1

in which:

P1	=	adjustment amount payable to the Supplier.

P0	=	Contract Price (base price).

a	=	fixed element representing profits and overheads included in the Contract Price and generally in the range of five (5) to fifteen (15) percent.

b	=	estimated percentage of labor component in the Contract Price.

c	=	estimated percentage of material component in the Contract Price.

L0, L1	=	*labor indices applicable to the appropriate industry in the country of origin on the base date and date for adjustment, respectively.

M0, M1	=	*material indices for the major raw material on the base date and date for adjustment, respectively, in the country of origin.

The Bidder shall indicate the source of the indices and the base date indices in its Bid.

The coefficients a, b, and c as specified by the Purchaser are as follows:

a = [insert value of coefficient]

b= [insert value of coefficient]

c= [insert value of coefficient]

Base date = thirty (30) days prior to the deadline for submission of the Bids.

Date of adjustment = [insert number of weeks] weeks prior to date of shipment (representing the mid-point of the period of manufacture).

The above price adjustment formula shall be invoked by either party subject to the following further conditions:

(a)	No price adjustment shall be allowed beyond the original delivery dates. As a rule, no price adjustment shall be allowed for periods of delay for which the Supplier is entirely responsible. The Purchaser will, however, be entitled to any decrease in the prices of the Goods and Services subject to adjustment.

(b)	If the currency in which the Contract Price P0 is expressed is different from the currency of origin of the labor and material indices, a correction factor will be applied to avoid incorrect adjustments of the Contract Price. The correction factor shall be: Z0 / Z1, where,

Z0  =  the number of units of currency of the origin of the indices which equal to one unit of the currency of the Contract Price P0 on the Base date, and

Z1  =  the number of units of currency of the origin of the indices which equal to one unit of the currency of the Contract Price P0 on the Date of Adjustment.

(c)	No price adjustment shall be payable on the portion of the Contract Price paid to the Supplier as advance payment.

Section X - Contract Forms

Table of Forms

Notification of Intention to Award	142

Beneficial Ownership Disclosure Form………………………………………………… 147

Letter of Acceptance	148

Contract Agreement	149

Performance Security	151

Advance Payment Security	155

Notification of Intention to Award

[This Notification of Intention to Award shall be sent to each Bidder that submitted a Bid, unless the Bidder has previously received notice of exclusion from the process at an interim stage of the procurement process.]

[Send this Notification to the Bidder’s Authorized Representative named in the Bidder Information Form]

For the attention of Bidder’s Authorized Representative

Name: [insert Authorized Representative’s name]

Address: [insert Authorized Representative’s Address]

Telephone/Fax numbers: [insert Authorized Representative’s telephone/fax numbers]

Email Address: [insert Authorized Representative’s email address]

[IMPORTANT: insert the date that this Notification is transmitted to Bidders. The Notification must be sent to all Bidders simultaneously. This means on the same date and as close to the same time as possible.]

DATE OF TRANSMISSION: This Notification is sent by: [email/fax] on [date] (local time)

Notification of Intention to Award

Purchaser: [insert the name of the Purchaser]

Project: [insert name of project]

Contract title: [insert the name of the contract]

Country: [insert country where RFB is issued]

Loan No. /Credit No. / Grant No.: [insert reference number for loan/credit/grant]

RFB No: [insert RFB reference number from Procurement Plan]

This Notification of Intention to Award (Notification) notifies you of our decision to award the above contract. The transmission of this Notification begins the Standstill Period. During the Standstill Period you may:

- request a debriefing in relation to the evaluation of your Bid, and/or

- submit a Procurement-related Complaint in relation to the decision to award the contract.

- The successful Bidder

| Name: | [insert name of successful Bidder] |
| --- | --- |
| Address: | [insert address of the successful Bidder] |
| Contract price: | [insert contract price of the successful Bid] |
| Total combined score: | [insert the total combined score of the successful Bidder] |

- Other Bidders [INSTRUCTIONS: insert names of all Bidders that submitted a Bid, Bid prices as read out and evaluated, technical scores and  combined scores.]

| Name of Bidder | Technical Score | Bid Price | Evaluated Bid Cost | Combined Score |
| --- | --- | --- | --- | --- |
| [insert name] | [insert Technical score] | [insert Bid price] | [insert evaluated cost] | [insert combined score] |
| [insert name] | [insert Technical score] | [insert Bid price] | [insert evaluated cost] | [insert combined score] |
| [insert name] | [insert Technical score] | [insert Bid price] | [insert evaluated cost] | [insert combined score] |
| [insert name] | [insert Technical score] | [insert Bid price] | [insert evaluated cost] | [insert combined score] |
| [insert name] | [insert Technical score] | [insert Bid price] | [insert evaluated cost] | [insert combined score] |

- Reason/s why your Bid was unsuccessful [Delete if the combined score already reveals the reason]

[INSTRUCTIONS: State the reason/s why this Bidder’s Bid was unsuccessful. Do NOT include: (a) a point by point comparison with another Bidder’s Bid or (b) information that is marked confidential by the Bidder in its Bid.]

- How to request a debriefing

DEADLINE: The deadline to request a debriefing expires at midnight on [insert date] (local time). You may request a debriefing in relation to the results of the evaluation of your Bid. If you decide to request a debriefing your written request must be made within three (3) Business Days of receipt of this Notification of Intention to Award. Provide the contract name, reference number, name of the Bidder, contact details; and address the request for debriefing as follows: Attention: [insert full name of person, if applicable] Title/position: [insert title/position] Agency: [insert name of Purchaser] Email address: [insert email address] Fax number: [insert fax number] delete if not used If your request for a debriefing is received within the 3 Business Days deadline, we will provide the debriefing within five (5) Business Days of receipt of your request. If we are unable to provide the debriefing within this period, the Standstill Period shall be extended by five (5) Business Days after the date that the debriefing is provided. If this happens, we will notify you and confirm the date that the extended Standstill Period will end. The debriefing may be in writing, by phone, video conference call or in person. We shall promptly advise you in writing how the debriefing will take place and confirm the date and time. If the deadline to request a debriefing has expired, you may still request a debriefing. In this case, we will provide the debriefing as soon as practicable, and normally no later than fifteen (15) Business Days from the date of publication of the Contract Award Notice.

- How to make a complaint

Period: Procurement-related Complaint challenging the decision to award shall be submitted by midnight, [insert date] (local time). Provide the contract name, reference number, name of the Bidder, contact details; and address the Procurement-related Complaint as follows: Attention: [insert full name of person, if applicable] Title/position: [insert title/position] Agency: [insert name of Purchaser] Email address: [insert email address] Fax number: [insert fax number] delete if not used At this point in the procurement process, you may submit a Procurement-related Complaint challenging the decision to award the contract. You do not need to have requested, or received, a debriefing before making this complaint. Your complaint must be submitted within the Standstill Period and received by us before the Standstill Period ends. Further information: For more information see the Procurement Regulations for IPF Borrowers (Procurement Regulations) (Annex III). You should read these provisions before preparing and submitting your complaint. In addition, the World Bank’s Guidance “How to make a Procurement-related Complaint” provides a useful explanation of the process, as well as a sample letter of complaint. In summary, there are four essential requirements: - You must be an ‘interested party’. In this case, that means a Bidder who submitted a Bid in this bidding process, and is the recipient of a Notification of Intention to Award. - The complaint can only challenge the decision to award the contract. - You must submit the complaint within the period stated above. - You must include, in your complaint, all of the information required by the Procurement Regulations (as described in Annex III).

- Standstill Period

DEADLINE: The Standstill Period is due to end at midnight on [insert date] (local time). The Standstill Period lasts ten (10) Business Days after the date of transmission of this Notification of Intention to Award. The Standstill Period may be extended as stated in Section 4 above.

If you have any questions regarding this Notification please do not hesitate to contact us.

On behalf of the Purchaser:

Signature: 	______________________________________________

Name:	______________________________________________

Title/position:	______________________________________________

Telephone:	______________________________________________

Email:		______________________________________________

INSTRUCTIONS TO BIDDERS: DELETE THIS BOX ONCE YOU HAVE COMPLETED THE FORM

This Beneficial Ownership Disclosure Form (“Form”) is to be completed by the successful Bidder.  In case of joint venture, the Bidder must submit a separate Form for each member.  The beneficial ownership information to be submitted in this Form shall be current as of the date of its submission.

For the purposes of this Form, a Beneficial Owner of a Bidder is any natural person who ultimately owns or controls the Bidder by meeting one or more of the following conditions:

- directly or indirectly holding 25% or more of the shares

- directly or indirectly holding 25% or more of the voting rights

- directly or indirectly having the right to appoint a majority of the board of directors or equivalent governing body of the Bidder

Beneficial Ownership Disclosure Form

RFB No.: [insert number of RFB process]

Request for Bid No.: [insert identification]

To: [insert complete name of Purchaser]

In response to your request in the Letter of Acceptance dated [insert date of letter of Acceptance] to furnish additional information on beneficial ownership: [select one option as applicable and delete the options that are not applicable]

(i) we hereby provide the following beneficial ownership information.

Details of beneficial ownership

| Identity of Beneficial Owner | Directly or indirectly holding 25% or more of the shares (Yes / No) | Directly or indirectly holding 25 % or more of the Voting Rights (Yes / No) | Directly or indirectly having the right to appoint a majority of the board of the directors or an equivalent governing body of the Bidder (Yes / No) |
| --- | --- | --- | --- |
| [include full name (last, middle, first), nationality, country of residence] |  |  |  |

OR

(ii) We declare that there is no Beneficial Owner meeting one or more of the following conditions:

- directly or indirectly holding 25% or more of the shares

- directly or indirectly holding 25% or more of the voting rights

- directly or indirectly having the right to appoint a majority of the board of directors or equivalent governing body of the Bidder

OR

(iii) We declare that we are unable to identify any Beneficial Owner meeting one or more of the following conditions. [If this option is selected, the Bidder shall provide explanation on why it is unable to identify any Beneficial Owner]

- directly or indirectly holding 25% or more of the shares

- directly or indirectly holding 25% or more of the voting rights

- directly or indirectly having the right to appoint a majority of the board of directors or equivalent governing body of the Bidder]”

Name of the Bidder: *[insert complete name of the Bidder]_________

Name of the person duly authorized to sign the Bid on behalf of the Bidder: **[insert complete name of person duly authorized to sign the Bid]___________

Title of the person signing the Bid: [insert complete title of the person signing the Bid]______

Signature of the person named above: [insert signature of person whose name and capacity are shown above]_____

Date signed [insert date of signing] day of [insert month], [insert year]_____

* In the case of the Bid submitted by a Joint Venture specify the name of the Joint Venture as Bidder. In the event that the Bidder is a joint venture, each reference to “Bidder” in the Beneficial Ownership Disclosure Form (including this Introduction thereto) shall be read to refer to the joint venture member.

** Person signing the Bid shall have the power of attorney given by the Bidder. The power of attorney shall be attached with the Bid Schedules.

Letter of Acceptance

[letterhead paper of the Purchaser]

[date]

To: [name and address of the Supplier]

Subject: Notification of award Contract No. . . . . . . . . . .

This is to notify you that your Bid dated . . . . [insert date] . . . . for execution of the . . . .  . . . . [insert name of the contract and identification number, as given in the SCC] . . . . . . . . . . for the Accepted Contract Amount of . . . . . . . . . [insert amount in numbers and words and name of currency], as corrected and modified in accordance with the Instructions to Bidders is hereby accepted by our Agency.

You are requested to furnish (i) the Performance Security within 28 days in accordance with the Conditions of Contract, using for that purpose one of the Performance Security Forms and (ii) the additional information on beneficial ownership in accordance with ITB 48.1 within eight (8) Business days using the Beneficial Ownership Disclosure Form, included in Section X, - Contract Forms, of the Bidding Document.

Authorized Signature:

Name and Title of Signatory:

Name of Agency:

Attachment: Contract Agreement

Contract Agreement

[The successful Bidder shall fill in this form in accordance with the instructions indicated]

THIS AGREEMENT made the [ insert: number ] day of [ insert: month ], [ insert: year ].

BETWEEN

(1)	[ insert complete name of Purchaser ], a [ insert description of type of legal entity, for example, an agency of the Ministry of .... of the Government of { insert name of Country of Purchaser }, or corporation incorporated under the laws of { insert name of Country of Purchaser } ] and having its principal place of business at [ insert address of Purchaser ] (hereinafter called “the Purchaser”), of the one part, and

(2)	[ insert name of Supplier ], a corporation incorporated under the laws of [ insert: country of Supplier ] and having its principal place of business at [ insert: address of Supplier ] (hereinafter called “the Supplier”), of the other part:

WHEREAS the Purchaser invited Bids for certain Goods and ancillary services, viz., [insert brief description of Goods and Services] and has accepted a Bid by the Supplier for the supply of those Goods and Services

The Purchaser and the Supplier agree as follows:

1.	In this Agreement words and expressions shall have the same meanings as are respectively assigned to them in the Contract documents referred to.

2.	The following documents shall be deemed to form and be read and construed as part of this Agreement. This Agreement shall prevail over all other contract documents.

- the Letter of Acceptance

- Letter of Bid - Technical Part

- Letter of Bid - Financial Part

- the Addenda Nos._____ (if any)

- Special Conditions of Contract

- General Conditions of Contract

- the Specification (including Schedule of Requirements and Technical Specifications)

- the completed Schedules (including Price Schedules)

- any other document listed in GCC as forming part of the Contract

3.	In consideration of the payments to be made by the Purchaser to the Supplier as specified in this Agreement, the Supplier hereby covenants with the Purchaser to provide the Goods and Services and to remedy defects therein in conformity in all respects with the provisions of the Contract.

4.	The Purchaser hereby covenants to pay the Supplier in consideration of the provision of the Goods and Services and the remedying of defects therein, the Contract Price or such other sum as may become payable under the provisions of the Contract at the times and in the manner prescribed by the Contract.

IN WITNESS whereof the parties hereto have caused this Agreement to be executed in accordance with the laws of [insert the name of the Contract governing law country] on the day, month and year indicated above.

For and on behalf of the Purchaser:

Signed: [insert signature]

in the capacity of [insert title or other appropriate designation]

in the presence of [insert identification of official witness]

For and on behalf of the Supplier:

Signed: [insert signature of authorized representative(s) of the Supplier]

in the capacity of [insert title or other appropriate designation]

in the presence of [insert identification of official witness]

Performance Security

Option 1: (Bank Guarantee)

[The bank, as requested by the successful Bidder, shall fill in this form in accordance with the instructions indicated]

[Guarantor letterhead or SWIFT identifier code]

Beneficiary: [insert name and Address of Purchaser]

Date: [Insert date of issue]

PERFORMANCE GUARANTEE No.: [Insert guarantee reference number]

Guarantor: [Insert name and address of place of issue, unless indicated in the letterhead]

We have been informed that _ [insert name of Supplier, which in the case of a joint venture shall be the name of the joint venture] (hereinafter called "the Applicant") has entered into Contract No. [insert reference number of the contract] dated [insert date] with the Beneficiary, for the supply of _ [insert name of contract and brief description of Goods and related Services] (hereinafter called "the Contract").

Furthermore, we understand that, according to the conditions of the Contract, a performance guarantee is required.

At the request of the Applicant, we as Guarantor, hereby irrevocably undertake to pay the Beneficiary any sum or sums not exceeding in total an amount of [insert amount in figures] 
(          ) [insert amount in words],1 such sum being payable in the types and proportions of currencies in which the Contract Price is payable, upon receipt by us of the Beneficiary’s complying demand supported by the Beneficiary’s statement, whether in the demand itself or in a separate signed document accompanying or identifying the demand, stating that the Applicant is in breach of its obligation(s) under the Contract, without the Beneficiary needing to prove or to show grounds for your demand or the sum specified therein.

This guarantee shall expire, no later than the …. Day of ……, 2…2, and any demand for payment under it must be received by us at this office indicated above on or before that date.

This guarantee is subject to the Uniform Rules for Demand Guarantees (URDG) 2010 Revision, ICC Publication No. 758, except that the supporting statement under Article 15(a) is hereby excluded.

_____________________ 
[signature(s)]

Note: All italicized text (including footnotes) is for use in preparing this form and shall be deleted from the final product.

Option 2: Performance Bond

By this Bond [insert name of Principal] as Principal (hereinafter called “the Supplier”) and [insert name of Surety] as Surety (hereinafter called “the Surety”), are held and firmly bound unto [insert name of Purchaser] as Obligee (hereinafter called “the Supplier”) in the amount of [insert amount in words and figures], for the payment of which sum well and truly to be made in the types and proportions of currencies in which the Contract Price is payable, the Supplier and the Surety bind themselves, their heirs, executors, administrators, successors and assigns, jointly and severally, firmly by these presents.

WHEREAS the Supplier has entered into a written Agreement with the Purchaser dated the 	 day of 	, 20 	, for [name of contract and brief description of Goods and related Services] in accordance with the documents, plans, specifications, and amendments thereto, which to the extent herein provided for, are by reference made part hereof and are hereinafter referred to as the Contract.

NOW, THEREFORE, the Condition of this Obligation is such that, if the Supplier shall promptly and faithfully perform the said Contract (including any amendments thereto), then this obligation shall be null and void; otherwise, it shall remain in full force and effect. Whenever the Supplier shall be, and declared by the Purchaser to be, in default under the Contract, the Purchaser having performed the Purchaser’s obligations thereunder, the Surety may promptly remedy the default, or shall promptly:

(1)	complete the Contract in accordance with its terms and conditions; or

(2)	obtain a Bid or Bids from qualified Bidders for submission to the Purchaser for completing the Contract in accordance with its terms and conditions, and upon determination by the Purchaser and the Surety of the lowest responsive Bidder, arrange for a Contract between such Bidder and Purchaser and make available as work progresses (even though there should be a default or a succession of defaults under the Contract or Contracts of completion arranged under this paragraph) sufficient funds to pay the cost of completion less the Balance of the Contract Price; but not exceeding, including other costs and damages for which the Surety may be liable hereunder, the amount set forth in the first paragraph hereof. The term “Balance of the Contract Price,” as used in this paragraph, shall mean the total amount payable by Purchaser to Supplier under the Contract, less the amount properly paid by Purchaser to the Supplier; or

(3)	pay the Purchaser the amount required by Purchaser to complete the Contract in accordance with its terms and conditions up to a total not exceeding the amount of this Bond.

The Surety shall not be liable for a greater sum than the specified penalty of this Bond.

Any suit under this Bond must be instituted not later than twenty-eight (28) days following the date of completion of the Supplier’s performance of its obligations under the Contract, including any warranty obligations.

No right of action shall accrue on this Bond to or for the use of any person or corporation other than the Purchaser named herein or the heirs, executors, administrators, successors, and assigns of the Purchaser.

In testimony whereof, the Supplier has hereunto set his hand and affixed his seal, and the Surety has caused these presents to be sealed with his corporate seal duly attested by the signature of his legal representative, this 	 day of 	 20 	.

SIGNED ON 	 on behalf of

By 	 in the capacity of

In the presence of

SIGNED ON 	 on behalf of

By 	 in the capacity of

In the presence of

Advance Payment Security

Demand Guarantee

[Guarantor letterhead or SWIFT identifier code]

Beneficiary: [Insert name and Address of Purchaser]

Date:	[Insert date of issue]

ADVANCE PAYMENT GUARANTEE No.:	[Insert guarantee reference number]

Guarantor: [Insert name and address of place of issue, unless indicated in the letterhead]

We have been informed that [insert name of Supplier, which in the case of a joint venture shall be the name of the joint venture] (hereinafter called “the Applicant”) has entered into Contract No. [insert reference number of the contract] dated [insert date] with the Beneficiary, for the execution of [insert name of contract and brief description of Goods and related Services] (hereinafter called "the Contract").

Furthermore, we understand that, according to the conditions of the Contract, an advance payment in the sum [insert amount in figures] () [insert amount in words] is to be made against an advance payment guarantee.

At the request of the Applicant, we as Guarantor, hereby irrevocably undertake to pay the Beneficiary any sum or sums not exceeding in total an amount of [insert amount in figures] 
(          ) [insert amount in words]1 upon receipt by us of the Beneficiary’s complying demand supported by the Beneficiary’s statement, whether in the demand itself or in a separate signed document accompanying or identifying the demand, stating either that the Applicant:

    - has used the advance payment for purposes other than toward delivery of Goods; or

    - has failed to repay the advance payment in accordance with the Contract conditions, specifying the amount which the Applicant has failed to repay.

A demand under this guarantee may be presented as from the presentation to the Guarantor of a certificate from the Beneficiary’s bank stating that the advance payment referred to above has been credited to the Applicant on its account number [insert number] at [insert name and address of Applicant’s bank].

The maximum amount of this guarantee shall be progressively reduced by the amount of the advance payment repaid by the Applicant as specified in copies of interim statements or payment certificates which shall be presented to us. This guarantee shall expire, at the latest, upon our receipt of a copy of the interim payment certificate indicating that ninety (90) percent of the Accepted Contract Amount, has been certified for payment, or on the [insert day] day of [insert month], 2 [insert year], whichever is earlier. Consequently, any demand for payment under this guarantee must be received by us at this office on or before that date.

This guarantee is subject to the Uniform Rules for Demand Guarantees (URDG) 2010 Revision, ICC Publication No.758, except that the supporting statement under Article 15(a) is hereby excluded.

____________________ 
[signature(s)]

Note: All italicized text (including footnotes) is for use in preparing this form and shall be deleted from the final product.

Non-performance, as decided by the Purchaser, shall include all contracts where (a) non-performance was not challenged by the contractor, including through referral to the dispute resolution mechanism under the respective contract, and (b) contracts that were so challenged but fully settled against the contractor. Non-performance shall not include contracts where Purchaser decision was overruled by the dispute resolution mechanism. Non-performance must be based on all information on fully settled disputes or litigation, i.e., dispute or litigation that has been resolved in accordance with the dispute resolution mechanism under the respective contract and where all appeal instances available to the applicant have been exhausted.

This requirement also applies to contracts executed by the Applicant as JV member.

The Proposer shall provide accurate information on the related Letter of Proposal about any litigation or arbitration resulting from contracts completed or ongoing under its execution over the last five years. A consistent history of awards against the Proposer or any member of a joint venture may result in failure of the Proposal.

For contracts under which the Proposer participated as a joint venture member or sub-contractor, only the Proposer’s share, by value, and role and responsibilities shall be considered to meet this requirement.

The amount of the Bond shall be denominated in the currency of the Purchaser’s country or the equivalent amount in a freely convertible currency.

For the avoidance of doubt, a sanctioned party’s ineligibility to be awarded a contract shall include, without limitation, (i) applying for pre-qualification, expressing interest in a consultancy, and bidding, either directly or as a nominated sub-contractor, nominated consultant, nominated manufacturer or supplier, or nominated service provider, in respect of such contract, and (ii) entering into an addendum or amendment introducing a material modification to any existing contract.

A nominated sub-contractor, nominated consultant, nominated manufacturer or supplier, or nominated service provider (different names are used depending on the particular bidding document) is one which has been: (i) included by the bidder in its pre-qualification application or bid because it brings specific and critical experience and know-how that allow the bidder to meet the qualification requirements for the particular bid; or (ii) appointed by the Borrower.

Inspections in this context usually are investigative (i.e., forensic) in nature.  They involve fact-finding activities undertaken by the Bank or persons appointed by the Bank to address specific matters related to investigations/audits, such as evaluating the veracity of an allegation of possible Fraud and Corruption, through the appropriate mechanisms.  Such activity includes but is not limited to: accessing and examining a firm's or individual's financial records and information, and making copies thereof as relevant; accessing and examining any other documents, data and information (whether in hard copy or electronic format) deemed relevant for the investigation/audit, and making copies thereof as relevant; interviewing staff and other relevant individuals; performing physical inspections and site visits; and obtaining third party verification of information.

For the avoidance of doubt, a sanctioned party’s ineligibility to be awarded a contract shall include, without limitation, (i) applying for pre-qualification, expressing interest in a consultancy, and bidding, either directly or as a nominated sub-contractor, nominated consultant, nominated manufacturer or supplier, or nominated service provider, in respect of such contract, and (ii) entering into an addendum or amendment introducing a material modification to any existing contract.

A nominated sub-contractor, nominated consultant, nominated manufacturer or supplier, or nominated service provider (different names are used depending on the particular bidding document) is one which has been: (i) included by the bidder in its pre-qualification application or bid because it brings specific and critical experience and know-how that allow the bidder to meet the qualification requirements for the particular bid; or (ii) appointed by the Borrower.

Inspections in this context usually are investigative (i.e., forensic) in nature.  They involve fact-finding activities undertaken by the Bank or persons appointed by the Bank to address specific matters related to investigations/audits, such as evaluating the veracity of an allegation of possible Fraud and Corruption, through the appropriate mechanisms.  Such activity includes but is not limited to: accessing and examining a firm's or individual's financial records and information, and making copies thereof as relevant; accessing and examining any other documents, data and information (whether in hard copy or electronic format) deemed relevant for the investigation/audit, and making copies thereof as relevant; interviewing staff and other relevant individuals; performing physical inspections and site visits; and obtaining third party verification of information.

1	 The Guarantor shall insert an amount representing the percentage of the Accepted Contract Amount specified in the Letter of Acceptance, and denominated either in the currency (ies) of the Contract or a freely convertible currency acceptable to the Beneficiary.

2	Insert the date twenty-eight days after the expected completion date as described in GC Clause 18.4. The Purchaser should note that in the event of an extension of this date for completion of the Contract, the Purchaser would need to request an extension of this guarantee from the Guarantor. Such request must be in writing and must be made prior to the expiration date established in the guarantee. In preparing this guarantee, the Purchaser might consider adding the following text to the form, at the end of the penultimate paragraph: “The Guarantor agrees to a one-time extension of this guarantee for a period not to exceed [six months] [one year], in response to the Beneficiary’s written request for such extension, such request to be presented to the Guarantor before the expiry of the guarantee.”

1	The Guarantor shall insert an amount representing the amount of the advance payment and denominated either in the currency(ies) of the advance payment as specified in the Contract, or in a freely convertible currency acceptable to the Purchaser.